## Sheet1
RMA | SPC Code | Part Number | S/N | Description | Fault Comments | Resolution Comments | Sender | Estimate Complete Time | Estimate Complete | Estimate Approved | Estimate Approved Time | QA Approved | QA Approved Time | Shipped | Shipped Time
| PSI | 32000-4 | TX19326C | CUST-TX. PSI. Error 15 | Error 5 and Error 15 stuck . No power when firing . Laser head require eval. ct | Reused all components. Realigned the laser head to MFG spec. YO | BCTechSupport@iridex.com | 2024-05-20 08:15 | Yes | No |  | No |  | No
| PSD | 30903-H500 | 1354776 | CUST-LIO. PSD. Dropped. Damaged knob & housing | Cannot test due to flexure mirror with delivery mirror dislocated, bent control |  | BCTechSupport@iridex.com | 2024-05-20 08:42 | Yes | No |  | No |  | No
| SD | 66120 | IQHG81008C | CUST-IQHG, WD, LOW POWER WITH FLEX FIBER, JUST SERVICE | CUSTOMER COMPLAINT LOW POWER, TESTED WITH IN HOUSE TOOL | TESTED WITH IN HOUSE TOOL WITH IN SPEC.JS | BCTechSupport@iridex.com | 2024-05-20 08:57 | Yes | No |  | No |  | No
| PSD | 66245-1 | IQI93435C | CUST-G6. PSD. PM. | Low power output. Need to realign laser head SD | Realign laser head and retested all the delivery devices. | BCTechSupport@iridex.com | 2024-05-20 09:05 | Yes | No |  | No |  | No
| PSD | 30903-H500 | K1831 | CUST-LIO. PSD. Broken fiber | Cannot test due to Twisted and burnt assembly fiber optic, burnt halogen bulb |  | BCTechSupport@iridex.com | 2024-05-20 09:48 | Yes | No |  | No |  | No
| SCD | 31172-01 | GH12661LM | CUST-SLA. SCD-P. Cracked ESF | CRACKED ESF AND BURNED FIBER. |  | BCTechSupport@iridex.com | 2024-05-20 11:18 | Yes | No |  | No |  | No
| PSI | 70292D | YEZ14948TC-D | CUST-TxCell. PSI. E09501=mirror motion error | E09501: Mirror Motion Error, Burned Fiber, Down Rev Rear PCBA, |  | BCTechSupport@iridex.com | 2024-05-20 12:43 | Yes | No |  | No |  | No
| PSI | 70292D | YEZ14948TC-D | CUST-TxCell. PSI. E09501=mirror motion error | E09501: Mirror Motion Error, Burned Fiber, Down Rev Rear PCBA, |  | BCTechSupport@iridex.com | 2024-05-20 12:45 | Yes | No |  | No |  | No
| PSD | 11180-H500 | L2049 | CUST-LIO. PSD. Broken knob & white light moves by itself | Cannot test due to twisted and burnt assembly fiber optic, broken control stem, |  | BCTechSupport@iridex.com | 2024-05-20 13:48 | Yes | No |  | No |  | No
| PSD | 13030 | 27972-LS | CUST-SLx. PSD. PM. Display will not light up sometime | LOW POWER HEAD SENT TO HR.JS |  | BCTechSupport@iridex.com | 2024-05-21 08:53 | Yes | No |  | No |  | No
| CEM | 13030D-220 | 28689-LS-D | CUST-SLx. CEM. Calibration due | 8888 ACROSS. INTERMITTENT. RY | Rplaced LCM, display board. Upgrade to 4.18. RY | BCTechSupport@iridex.com | 2024-05-21 08:50 | Yes | No |  | No |  | No
| SD | 65600-08 | IQY72653C | CUST-IQ577. WD. E05111. Intermittently recognizes LIO | HAVE NOT SEEN CUSTOMER COMPLAINT FIBER WORKS MAY ASK TO |  | BCTechSupport@iridex.com | 2024-05-21 08:16 | Yes | No |  | No |  | No
| SI | 70292 | YEZ13694TC | CUST-Txcell. PSI. E09500 | UNABLE TO TEST GALVO'S ARE BAD CAUSING THE ERROR E09500 | REPLACED GALVO'S INTERFACE BOARD FIBER USB CABLE'S | BCTechSupport@iridex.com | 2024-05-22 09:20 | Yes | No |  | No |  | No
SRV-RMA00979 | WD | SA-06423 |  | Warranty |  |  | BCTechSupport@iridex.com | 2024-05-22 10:27 | Yes | No |  | No |  | No
| PSD | 70295 | YH12886TC | CUST-TxCell. PSD. LED blinking non-stop | Blinking Non-Stop. Clean Lenses, Mirrors, and Test | Could not Duplicate Problem. Cleaned Lenses, Mirrors and Tested. | BCTechSupport@iridex.com | 2024-05-23 06:59 | Yes | No |  | No |  | No
| PSD | 70294 | YEC13043TC | CUST-TxCell, EO5111 PERSISTS EVEN CHANGE | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING,COMPUTER BOARD NOT WORKIN. |  | BCTechSupport@iridex.com | 2024-05-24 10:44 | Yes | No |  | No |  | No
| SI | 66245-1 | IQI92848C | CUST-G6. PSI. 15% low power | LOW POWER HEAD NEED REPAIR.JS | Realign and calibrationlaser head. | BCTechSupport@iridex.com | 2024-05-28 06:29 | Yes | No |  | No |  | No
| SI | 13030 | 30257-LS | CUST-SLx. WI. No illumination light. Use REV N p/s.. Repaired 3 months ago |  |  | BCTechSupport@iridex.com | 2024-05-28 07:00 | Yes | No |  | No |  | No
| WD | 31703 | 100636F | 31703, W/R, RETURN TO INVENTORY, INTERMITTENTLY STOP WORKING | Red light flashing. No function. |  | BCTechSupport@iridex.com | 2024-05-28 08:51 | Yes | No |  | No |  | No
| PSD | 13030 | 30176-LS | CUST-SLx, PSD, PM CHECKUP |  |  | BCTechSupport@iridex.com | 2024-05-28 09:03 | Yes | No |  | No |  | No
| PSD | 70292 | YEZ14136TC | CUST-TxCell. PSI. Damaged "Y" adjust/mirror | BROKEN BEZEL AND SPRING TORSION IS DAMAGE SOFTWARE DOWN REV. | REPLACED BEZEL, AND SPRING TORSION. REPLACED DOWN REV SOFTWARE | BCTechSupport@iridex.com | 2024-05-29 06:57 | Yes | No |  | No |  | No
| PSD | 32000-4 | TX19382C | CUST-TX. PSD. Fiber adapter is not centered. Display flashes when LIO is connected. Bad display pcb | Fiber connector is not on the center. Retest no problem found from the customer | Adjust fiber adapter .Retest all delivery devices and burn in ok . | BCTechSupport@iridex.com | 2024-05-29 10:53 | Yes | No |  | No |  | No
| PSI | 70295D | YH14212TC-D | CUST-Txcell. PSI. E05111 intermitttently | unable to tested control box non-stop blinking and has an error massage. aiming | replaced interface board and usb cable,& bezel, tested aligned calibrated clean | BCTechSupport@iridex.com | 2024-05-30 07:53 | Yes | No |  | No |  | No
| WI | 70295 | YH16472TC | CUST-TXCell. WI. OBF! No aiming beam. fiber may be broken. Tested laser with endo and worked fine |  |  | BCTechSupport@iridex.com | 2024-05-30 08:17 | Yes | No |  | No |  | No
| WD | 70295 | YH16188TC | CUST-TXCELL. WD. OBF, LED KEEP BLINKING, JUST INSTALL AND UNIT FAILD | LED LIGHT BLINKING DUE TO COMPUTER NOT BOOT UP,BROKEN FIBER. |  | BCTechSupport@iridex.com | 2024-05-30 10:01 | Yes | No |  | No |  | No
| PSI | 70295 | YH15583TC | CUST-TxCell. PSI. LED blinking non-stop | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING COMPUTER BOARD IS BAD.INTERFCACE |  | BCTechSupport@iridex.com | 2024-05-30 11:02 | Yes | No |  | No |  | No
| PSD | 70295 | YH13132TC | CUST-Txcell. PSD. LED blinking non-stop | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING AND COMPUTER BOARD OVER HEATED |  | BCTechSupport@iridex.com | 2024-05-31 06:22 | Yes | No |  | No |  | No
| WD | 70292 | YEZ16317TC | OBF-CUST.WI. OBF Per Brian, cannot center the beam with the "X"lever. it keeps on drifting to the R | Confirm same issue found as reported by customer. Unit originally a demo return | Found (2) loose screws on the Lever Assy P/N: 31039, which was causing | BCTechSupport@iridex.com | 2024-06-03 07:35 | Yes | No |  | No |  | No
| SD | 12733 | 2802-VET | CUST-VET. SCD-G. No illumination light. Changed bulb. bad power supply | CUSTOMER COMPLAINT NO ILLUMINATION LIGHT.JS | NEED POWER SUPPLY.JS | BCTechSupport@iridex.com | 2024-06-04 09:37 | Yes | No |  | No |  | No
| PSD | 70292 | YEZ14619TC | CUST-TxCell. PSD. Fiber damaged/low pwoer output | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING AND FIBER IS BURNED. |  | BCTechSupport@iridex.com | 2024-06-05 10:18 | Yes | No |  | No |  | No
| PSD | 70295 | YH13167TC | CUST-TxCell,PSD, LEDblinking non stop | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING, BURNED FIBER |  | BCTechSupport@iridex.com | 2024-06-05 11:22 | Yes | No |  | No |  | No
| WD | 30903-H500 | K2010 | CUST-LIO. WD. Optics not aligned | Cannot test due to ocular housing dislocated, loose viewing assembly, missing | Replaced eyepiece rubber, Heine halogen bulb, adjusted acular housing, T-slot | BCTechSupport@iridex.com | 2024-06-05 14:04 | Yes | No |  | No |  | No
| PSD | 70067 | EZZ12927TC | CUST-TxCell. PSD. LED blinking non-stop | UNABLE TO FIXED. THE NEW COMPUTER BOARD IS NOT COMPATIBLE WITH THIS UNIT |  | BCTechSupport@iridex.com | 2024-06-06 06:49 | Yes | No |  | No |  | No
| PSD | 70057 | YEZ12472TC | CUST-TxCell. PSD. E05111 | UNABLE TO FIXED THE COMPUTER BOARD NOT WORKING AND THE NEW |  | BCTechSupport@iridex.com | 2024-06-06 06:52 | Yes | No |  | No |  | No
| PSD | 70057 | YEZ12736TC | CUST-TxCell. PSD. Broken fiber | UNABLE TO TESTED COMPUTER BOARD NOT WORKING |  | BCTechSupport@iridex.com | 2024-06-06 06:52 | Yes | No |  | No |  | No
|  | 12729D | 27383-LS-D | CUST-SLx. CEM. Calibration due | PM CHECK.JS | RETESTED ALL THE DELIVERY DEVICES.JS | BCTechSupport@iridex.com | 2024-06-11 05:42 | Yes | No |  | No |  | No
| PSD | 13152-H | 1010202 | CUST-LIO. PSD. Low power | Connot test due to twisted and burnt assembly fiber optic, burnt halogen bulb, | Replaced assembly fiber optic, control stem, eccentric shaft, halogen bulb, dust | BCTechSupport@iridex.com | 2024-06-11 11:52 | Yes | No |  | No |  | No
| SD | 65500 | IQG81583C | CUST-IQ532. PSD. Port 1 & 2 not detecting fiber |  |  | BCTechSupport@iridex.com | 2024-06-12 05:45 | Yes | No |  | No |  | No
| PSD | 30685 |  | CUST-SLA, PSD, CHECK SMARTKEY CABLE & ALIGNMENT |  |  | BCTechSupport@iridex.com | 2024-06-12 09:42 | Yes | No |  | No |  | No
| PSD | 12729 | 27277 | CUST-SLx. PSD. Low power output | DUPLICATED CUSTOMER COMPLAINT LOW POWER,HEAD SENT TO HEAD |  | BCTechSupport@iridex.com | 2024-06-12 09:47 | Yes | No |  | No |  | No
| PSD | 30903-H500 | 1344124 | CUST-LIO. PSD. Fiber damaged | Cannot test due to Twisted and burnt assembly fiber optic, wrong and burnt |  | BCTechSupport@iridex.com | 2024-06-12 10:31 | Yes | No |  | No |  | No
| SD | 65777 | RC0361 | CUST-REMOTE , PSD, POWERS JUMPS UP AND OWN UNABLE TO DIAL POWERS |  |  | BCTechSupport@iridex.com | 2024-06-12 10:57 | Yes | No |  | No |  | No
| SD | 65777 | RC10319 | CUST-REMOTE. PSD. Not working. Tested 2nd remote & worked fine | NOT TURN ON MAY HAS CABLE SHORT.NEED TO REPLACE CABLE.JS |  | BCTechSupport@iridex.com | 2024-06-12 13:13 | Yes | No |  | No |  | No
| PSD | 65900-H500 | K1752 | CIST-LIO. PSD. Mirror out of alignment | Cannot test due to twisted and broken assembly fiber optic, wrong and burnt |  | BCTechSupport@iridex.com | 2024-06-12 14:12 | Yes | No |  | No |  | No
| PSD | 70295 | YH14402TC | CUST-TxCell. PSD. Checkup/service. May have problem | SC500 CABLE IS BAD | REPLACED SC500 CABLE TESTED ALIGNED CALIBRATED CLEAN AND INSPECTED | BCTechSupport@iridex.com | 2024-06-13 05:21 | Yes | No |  | No |  | No
| SCD | 65622 | RYEZ16072 | CUST-SLA. SCD-P. Low power. Fiber burnt. | Burned fiber, | Replaced fiber, test, clean and re-aligment. | BCTechSupport@iridex.com | 2024-06-13 08:26 | Yes | No |  | No |  | No
| SCD | 30903-H | 1143424 | CUST-LIO,SCD,NO AIMBEAM MAY BE BROKEN FIBER | Cannot test due to twisted and broken assemby fiber optic, broken control stem, |  | BCTechSupport@iridex.com | 2024-06-13 08:45 | Yes | No |  | No |  | No
| PSD | 31172-02 | GH16090LM | CUST-SLA. PSD. Connect smartkey message | Burned fiber, | Replaced fiber, test,clean , re-aligment and inspect. | BCTechSupport@iridex.com | 2024-06-13 10:38 | Yes | No |  | No |  | No
| PSD | 30903-H | 1143462 | CUST-LIO, PSD, BROKEN KNOBS | Cannot test due to Twisted and burnt assembly fiber optic, broken control stem, |  | BCTechSupport@iridex.com | 2024-06-13 10:55 | Yes | No |  | No |  | No
| PSD | 30903-H500-L | 1351091-L | CUST-LIO. PSD. Fiber/knob damaged | Cannot test due to Twisted and broken assembly fiber optic, broken control stem, |  | BCTechSupport@iridex.com | 2024-06-13 12:05 | Yes | No |  | No |  | No
| SCD | 30903-H500 | 1361982 | CUST-LIO. SCD-PP. Broken knobs | Cannot test due to broken control stem, viewing assembly dislocated from T_slot |  | BCTechSupport@iridex.com | 2024-06-13 13:08 | Yes | No |  | No |  | No
| PSD | 31172-01 | GH9072LM | CUST-SLA, PSD, BROKEN SMART KEY, NO SPOT SIZE DETECTED | Damaged smartkey and fiber. | Replaced fiber, cable.Re-aligment, test,clean and inspect. | BCTechSupport@iridex.com | 2024-06-13 13:17 | Yes | No |  | No |  | No
| PSI | 66245-1 | IQI92273C | Cust-G6-PM/GUI-PSI $1250 flat rate | NOT LOADINGH APPLICATION, DAMAGED ENCODER NEED TO REPLACE.JS | REPLACED ENCODER AND RETESTED.JS | BCTechSupport@iridex.com | 2024-06-14 10:52 | Yes | No |  | No |  | No
| PSD | 65505 | REZZ14603 | CUST-SLA. PSD. Extreme pain during the treatment w/IQ532 | No problem found. | No problem found .Inspect and test, clean . | BCTechSupport@iridex.com | 2024-06-14 11:30 | Yes | No |  | No |  | No
| WD | 87300 | 1105071212 | CUST-LIO. WD. Light flickers on and off if you're not careful with the wire that connects it | Cannot test due to damaged white light switch cable will need recharge the | Replaced white light switch cable, adjsuted beam centered, recharged battery and | BCTechSupport@iridex.com | 2024-06-14 11:48 | Yes | No |  | No |  | No
| LD | 31172-01-L | GH13522LM-L | LNR,LD, BACK TO INVENTORY |  |  | BCTechSupport@iridex.com | 2024-06-14 13:25 | Yes | No |  | No |  | No
| PSI | 70292 | YEZ13671C | CUST-TxCell. PSI. E09025 | COMPUTER BOARD NOT WORKING AND ASSY, AIMING DIODE, TO |  | BCTechSupport@iridex.com | 2024-06-17 11:09 | Yes | No |  | No |  | No
| PSD | 12733 | 1115-VET | CUST-VET. PSD. Error 15. | Error 15. SD. |  | BCTechSupport@iridex.com | 2024-06-18 06:09 | Yes | No |  | No |  | No
| SI | 65600-08 | IQY72706C | CUST-IQ577. WI. OBF! UIM application failed to load message | UIM FAIL TOLOAD APPLICATION,AS CUSTOMER COMPLAINT.JS |  | BCTechSupport@iridex.com | 2024-06-18 09:42 | Yes | No |  | No |  | No
| SD | 65500 | IQG81532C | CUST-IQ532. SCD-PP. Halo light intermittently comes on | CUSTOMER COMPLAINT ILLUMINATION LIGHT INTERMITTENTLY WORK |  | BCTechSupport@iridex.com | 2024-06-18 10:31 | Yes | No |  | No |  | No
| PSI | 70292 | YEZ15989TC | CUST-TxCell. PSI. computer corrupted | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING. |  | BCTechSupport@iridex.com | 2024-06-19 06:11 | Yes | No |  | No |  | No
| PSD | 65511 | GH13949RM | CUST-SLA. PSD. Damaged smartkey connector | Damaged smartkey cable and burned fiber. |  | BCTechSupport@iridex.com | 2024-06-19 07:43 | Yes | No |  | No |  | No
| SD | 66120 | IQHG81010C | CUST-IQ532XP. SCD-PP. PM. No issue | UNIT CAME FOR PM FOUND LOW POWER HEAD SENT TO HEAD ROOM.JS |  | BCTechSupport@iridex.com | 2024-06-19 08:12 | Yes | No |  | No |  | No
| PSI | 70295 | YH14298TC | CUST-TxCell. PSI. Scanning aiming beam mount may be broken | BROKEN BEZEL AND BROKEN FLAX PLATE. AND PCAM BOARD NOT WORKING. |  | BCTechSupport@iridex.com | 2024-06-20 09:57 | Yes | No |  | No |  | No
| PSD | 65504 | GH14654RV | CUST-SLA. PSD. Dropped & damaged | Unit was drop . Damaged mirror assembly. | Replaced parts ,Aligment, test, clean and inspec. | BCTechSupport@iridex.com | 2024-06-20 10:29 | Yes | No |  | No |  | No
| DCI | 66245-1 | IQI91982C | DCI. Service & return G6 to Carleton. E05103/low power | Dead battery , ER6006 photocell are not reading match. | Replaced 1 optic (11586). Realign and calibration. RY. | BCTechSupport@iridex.com | 2024-06-20 13:29 | Yes | No |  | No |  | No
| WD | 31172-02 | GH16709LM | OBF. WD. Return inventory. Laser reset when fiber is touched |  |  | BCTechSupport@iridex.com | 2024-06-21 08:36 | Yes | No |  | No |  | No
| PSD | 30266 | GH7330 | CSUT-SLA. PSD. Broken/damage fiber |  |  | BCTechSupport@iridex.com | 2024-06-21 12:41 | Yes | No |  | No |  | No
| PSD | 30903-H500 | 1361499 | CUST-LIO. PSD. Fiber damaged | Cannot test due to Twisted and burnt assembly fiber optic, burnt halogn bulb, |  | BCTechSupport@iridex.com | 2024-06-21 15:53 | Yes | No |  | No |  | No
| PSD | 66245-1 | IQI92031C | CUST-CYCLO-G6,PSD,PROBE CONNECTOR STUCK IN THE PORT | Fiber adapter stuck ,can't firing .Found damaged fiber pcb #65143 and #10153 | Replaced fiber adapter 10153. Realign and calibration laser head. RY | BCTechSupport@iridex.com | 2024-06-24 08:16 | Yes | No |  | No |  | No
| PSD | 30903-H | 1251501 | CUST-LIO. PSD. Fiber damaged | Cannot test duet to Twisted. broken and burnt assembly fiber optic, bent |  | BCTechSupport@iridex.com | 2024-06-24 09:28 | Yes | No |  | No |  | No
| SI | 12729 | 27080 | CUST-SLx. PSI. Low power | LOW POWER AS CUSTOMER MENTION, HEAD SENT TO HEAD ROOM.JS |  | BCTechSupport@iridex.com | 2024-06-24 10:26 | Yes | No |  | No |  | No
| PSD | 65900-H500 | K1384 | CUST-LIO. PSD. Cannot focus | Cannot test due to twisted and burnt, missing eyepiece rubber, damaged extension |  | BCTechSupport@iridex.com | 2024-06-24 12:35 | Yes | No |  | No |  | No
| PSD | 66245 | IQI91039C | CUST-G6. PSD. UIM application failed to load | UIM applicatiom failed to load, broken encoder and power ouput low . SD | Replace broken encoder | BCTechSupport@iridex.com | 2024-06-25 06:11 | Yes | No |  | No |  | No
| PSD | 30898-Z30-L | EZZ7630-L | LOANER-SLA.LD.RETURN INTO INVENTORY |  |  | BCTechSupport@iridex.com | 2024-06-25 12:55 | Yes | No |  | No |  | No
| PSD | 66245 | IQI91275C | Cust-G6-PM-PSD-Not recognizing probes | Not recognizing probes. SD | Replace PCBA fiber detect and retested all the delivery devices. SD | BCTechSupport@iridex.com | 2024-06-25 14:31 | Yes | No |  | No |  | No
| SD | 65600-01 | IQY71753C | CUST-IQ577. SCD-PP. Pattern grid is not working properly | CUSTOMER COMPLAINT SCANNER NOT WORKING | RETESTED ALL THE DELIVERY DEVICES.JS | BCTechSupport@iridex.com | 2024-07-08 05:29 | Yes | No |  | No |  | No
| PSI | 70295 | YH14832TC | CUST-TxCell. PSI. E09500 | E09500 SCANNER CONTROLLER NOT FOUND, ERROR | REPLACED POWER SUPPLY, TESTED ALIGNED CALIBRATED CLEAN AND INSPEC TED. | BCTechSupport@iridex.com | 2024-07-08 06:18 | Yes | No |  | No |  | No
| SD | 65600-01 | IQY71942C | CUST-IQ577,PSD, UPON TURN ON BLACK SCREEN NO DISPLAY |  |  | BCTechSupport@iridex.com | 2024-07-08 10:12 | Yes | No |  | No |  | No
| SI | 65600-01 | IQY71863C | CUST-IQ577. PSI. E06010 message |  |  | BCTechSupport@iridex.com | 2024-07-08 11:13 | Yes | No |  | No |  | No
| WD | 66245 | IQI91151C | Cust-G6-GUI-WD | UNIT CAME FOR UPGRADE,WATING FOR UPGRADE TO FINAL TEST.JS |  | BCTechSupport@iridex.com | 2024-07-09 07:19 | Yes | No |  | No |  | No
| SD | 13030 | 28581-LS | CUST-SLx. PSD. Low power output | DUPLICATED CUST COMPLAINT LOW POWER. HEAD ROOM.JS |  | BCTechSupport@iridex.com | 2024-07-09 07:59 | Yes | No |  | No |  | No
| SD | 13030 | 29199-LS | CUST-SLx. PSD. PM. |  |  | BCTechSupport@iridex.com | 2024-07-10 09:59 | Yes | No |  | No |  | No
| PSD | 13030 | 29754-LS | CUST-SLx. PSD. low power output/no illumination light | no illumination light need to replace power supply.js | replaced power supply and retested. js | BCTechSupport@iridex.com | 2024-07-10 10:20 | Yes | No |  | No |  | No
| PSD | 30268-H500 | L2133 | CUST-LIO. PSD. Fiber damaged RMA changed from 01660 to | Cannot test due to twisted and burnt assembly fiber optic, burn halogen bulb, |  | BCTechSupport@iridex.com | 2024-07-10 15:32 | Yes | No |  | No |  | No
| SI | 13030-220 | 29625-LS | CUST-SLx. PSI. Error 16. No aiming beam. Counter reset button not working | NOT BOOTING UP DAMAGED COUNTER RESET AND TREAT SWITCH NEED TO |  | BCTechSupport@iridex.com | 2024-07-11 09:29 | Yes | No |  | No |  | No
| PSD | 30268-H | 1045046 | CUST-LIO,PSD,BROKEN &OUT OF AKLIGNMENT | Canno test due to lamp connector damage white light read low power 4mW, missing | Replaced control stem, eccentric shaft, dust cap with lanyard, eyepiece rubber, | BCTechSupport@iridex.com | 2024-07-15 15:43 | Yes | No |  | No |  | No
| PSD | 70297 | EZZ13810TC | CUST-TxCell. PSD. MM handle is not moving the aiming beam directionally | BURNED FIBER,BROKEN BEZEL NEED TO REPLACED INTERFACE BOARD. |  | BCTechSupport@iridex.com | 2024-07-16 05:58 | Yes | No |  | No |  | No
| PSD | 65900-H500 | 1362493 | CUST-LIO. PSD. No aiming beam. fiber broken | Cannot test due to broken assembly fiber optic, burnt halogen bulb, bent control |  | BCTechSupport@iridex.com | 2024-07-16 07:54 | Yes | No |  | No |  | No
| PSI | 30904 | 2394F | CUST-HEAD-GL,PSI,LPA | Need to replace diode. YO |  | BCTechSupport@iridex.com | 2024-07-16 08:42 | Yes | No |  | No |  | No
| PSI | 66245-1 | IQI92621C | CUST-CYCLO-G6,PSD,DEAD BATTERY |  | Replaced 2 cables 65215,65211 clean all the optics. Realign and calibration. | BCTechSupport@iridex.com | 2024-07-22 06:55 | Yes | No |  | No |  | No
| PSD | 66245 | IQI91213C | C UST-CYCLO-G6,PSD, ERROR E06002 | Low output power. Need to realign laser head. SD | Realign and calibration. | BCTechSupport@iridex.com | 2024-07-22 12:45 | Yes | No |  | No |  | No
| PSD | 70295 | YH15208TC | CUST-TxCell.PSD. PM | BURNED FIBER. |  | BCTechSupport@iridex.com | 2024-07-22 12:53 | Yes | No |  | No |  | No
| PSD | 30903-H500 | 1345597 | CUST-LIO, PSD, MIRROR FELL OFF | Cannot test due to damaged sma connector, damaged lamp connector white light |  | BCTechSupport@iridex.com | 2024-07-22 13:45 | Yes | No |  | No |  | No
| PSD | 30903-H | 1046520 | CUST-LIO, PSD, BROKEN KNOBS | Cannot test due to Twisted and burnt assembly fiber optic, broken control stem, |  | BCTechSupport@iridex.com | 2024-07-22 15:17 | Yes | No |  | No |  | No
| PSD | 65900-H500 | 1361097 | CUST-LIO, PSD, NO ILLUMINATION CHANGING ILLUM CABLE AND BULB | Cannot test due to Twisted and burnt assembly fiber optic, burnt Halogen bulb, |  | BCTechSupport@iridex.com | 2024-07-23 11:34 | Yes | No |  | No |  | No
| PSD | 30903-H500 | K1969 | CUST-LIO. PSD. Broken mirror | Cannot test due to Twisted and broken assembly fiber optic, Wrong and burnt bulb |  | BCTechSupport@iridex.com | 2024-07-23 11:47 | Yes | No |  | No |  | No
| PSD | 30903-H500 | K1012 | CUST-LIO. PSD. Aiming beam misaligned | Cannot test due to Twisted and burnt assembly fiber optic, missing eyepiece |  | BCTechSupport@iridex.com | 2024-07-23 12:59 | Yes | No |  | No |  | No
| PSD | 70295 | YH16205TC | CUST-TxCell, PSD, LED KEEP BLINKING & CHECK FIBER OPTIC CABLE | CRONTROL BOX NON-STOP BLINKING,MICRO MANIPULATER IS LOOSE, |  | BCTechSupport@iridex.com | 2024-07-23 13:17 | Yes | No |  | No |  | No
| SD | 65600-01 | IQY71170C | CUST-IQ577,PSD, JUMPY ENCODER | PORT ONE NOT RECOGNISE FIBER PORT TWO LOW POWER |  | BCTechSupport@iridex.com | 2024-07-24 09:52 | Yes | No |  | No |  | No
| PSI | 65600-01 | IQY71306C | CUST-IQ577. PSI. E06002-power output out of range | Power output out of range. SD |  | BCTechSupport@iridex.com | 2024-07-24 15:01 | Yes | No |  | No |  | No
| SD | 65500-08 | IQG82031C | Cust-532-WD-In for service & doesn't work still | AT START UP FAN FULL BLOW AND DISCONNECT LASER AND HAS BEEPING | -Replaced LCM s/n 1JVAN2X to 1JVAP0U realigned the laser head, recalibrated | BCTechSupport@iridex.com | 2024-07-25 06:32 | Yes | No |  | No |  | No
| PSD | 65512 | YH11433RM | CUST-SLA, PSD,BROKEN LENSE AND OUT OF ALIGNMENT | Burned fiber ,broken yoke assembly,loose screw. | Replaced parts. Test, clean and inspec. | BCTechSupport@iridex.com | 2024-07-25 10:18 | Yes | No |  | No |  | No
| PSD | 65900-H500 | 1353144 | CUST-LIO.PSD.PM | Cannote test due to Clippig aming beam, burnt halogen bulb, unit very dirty and |  | BCTechSupport@iridex.com | 2024-07-25 12:42 | Yes | No |  | No |  | No
| PSI | 30685 | GH7876M | CUST-SLA.PSI. PM |  |  | BCTechSupport@iridex.com | 2024-07-25 12:57 | Yes | No |  | No |  | No
| PSD | 70295 | YH15828TC | CUST-TxCell. PSD. LED blinking non-stop | CONTROL BOX NON-STOP BLINKING. | REPLACED COMPUTER BOARD TESTED ALIGNED CALIBRATED CLEAN AND INSPEC TED. | BCTechSupport@iridex.com | 2024-07-27 12:56 | Yes | No |  | No |  | No
| PSI | 65600-01 | IQY71365C | CUST-IQ577. PSI. E05111/Wavelength incompatible after replacing fiber detect pcb | DUPLICATED CUSTOMER COMPLAINT FIBER NOT RECOGNISE PORT | HEAD SENT TO HEAD ROOM UPGRADED FFC CABLE.JS | BCTechSupport@iridex.com | 2024-07-29 10:24 | Yes | No |  | No |  | No
SRV-RQ001896 | SD | 12731 | 2440 | CUST-SL, PSD, LOW POWER | LOW POWER AS CUSTOMER COMPLAINT. HEAD ROOM JS | Replaced aiming beam diode, adjust screws and realign laser head. RY | BCTechSupport@iridex.com | 2024-07-31 06:46 | Yes | No |  | No |  | No
SRV-RQ001690 | SI | 13030 | 28743-LS | CUST-SLx. PSI. Low power output | LOW POWER AS CUSTOMER SPECIFIED.H.R.JS |  | BCTechSupport@iridex.com | 2024-07-31 08:00 | Yes | No |  | No |  | No
SRV-RQ001796 | SI | 65600-08 | IQY72772C | CUST-IQ577. WI. No response & Beeping noise when touchscreen is depressed |  |  | BCTechSupport@iridex.com | 2024-08-02 10:09 | Yes | No |  | No |  | No
SRV-RQ001907 | PSD | 70297 | EZZ13823TC | CUST-TxCell, PSD, BROKEN, MAYBE LED FLASHING |  |  | BCTechSupport@iridex.com | 2024-08-05 11:07 | Yes | No |  | No |  | No
SRV-RQ001936 | SD | 65600-01 | IQY71344C | CUST IQ577. WD. PM | PM CHECK.JS | RETESTED ALL THE DELIVERY DEVICES.JS | BCTechSupport@iridex.com | 2024-08-06 08:56 | Yes | No |  | No |  | No
SRV-RQ001663 | WI | 70292 | YEZ16227TC | CUST-TxCell. WI. E05111. LED blinking non-stop | CONTROL BOX NON-STOP BLINKING COMPUTER BOARD IS BAD CAUSING THE PROBLEM. |  | BCTechSupport@iridex.com | 2024-08-06 10:36 | Yes | No |  | No |  | No
SRV-RQ001809 | PSD | 31172-01 | GH13802LM | CUST-SLA. PSD. Connect smartkey message/bad smartkey cable assy | Broken fiber and damaged cable. | Replaced cable , test, clean and inspec. | BCTechSupport@iridex.com | 2024-08-06 13:29 | Yes | No |  | No |  | No
SRV-RQ001903 | SCD | 31172-01 | GH11848LM | CUST-SLA, SCD,BROKEN FIBER , LOW POWER | No problem found, | Test, clean and inspec. | BCTechSupport@iridex.com | 2024-08-06 13:36 | Yes | No |  | No |  | No
SRV-RQ001858 | WD | 65515-H500 | K0814 | CUST-LIO. WD. Dim aiming beam/fiber damaged |  |  | BCTechSupport@iridex.com | 2024-08-07 11:35 | Yes | No |  | No |  | No
SRV-RQ001931 | SD | 66245 | IQI91710C | Cust-G6, PSD, PM and PM certification needed |  |  | BCTechSupport@iridex.com | 2024-08-08 10:22 | Yes | No |  | No |  | No
SRV-RQ001904 | SI | 65500 | IQG81641C | CUST-IQ532, WI, requesting a PM (Including a TxCell). |  |  | BCTechSupport@iridex.com | 2024-08-09 07:31 | Yes | No |  | No |  | No
SRV-RQ001861 | PSD | 31172-01 | GH15091LM | CUST-SLA, PSD, NO AIM BEAM NO POWER |  |  | BCTechSupport@iridex.com | 2024-08-14 07:07 | Yes | No |  | No |  | No
SRV-RQ001919 | PSD | 31172-01 | GH15384LM | CUST-SLA, PSD, FIBER PULLED OUT, DAMAGED FIBER |  |  | BCTechSupport@iridex.com | 2024-08-13 14:14 | Yes | No |  | No |  | No
SRV-RQ001930 | SD | 12733 | 2726-VET | CUST-VET, PSD, NO ILLUMINATION | DUPLICATED CUSTOMER COMPLAIT NO ILLUMINATION.JS | Replaecd adjust screws and realign laser head. RY | BCTechSupport@iridex.com | 2024-08-13 12:24 | Yes | No |  | No |  | No
SRV-RQ001791 | SI | 66245 | IQI91267C | CUST-G6. PSI. E06002 with MP3 but G-probe worked fine | CUSTOMER COMPLAINT ENT HAS PROBLEM FOUND UNIT NOT WORKING | REPLACED ALL THE CABLES AND RETESTED ALL THE DELIVERY DEVICES | BCTechSupport@iridex.com | 2024-08-13 09:03 | Yes | No |  | No |  | No
SRV-RQ001918 | PSD | 66245-1 | IQI92573C | Cust-G6-PSD-PM | Power ouput low | Realign and calibration Laser head. RY | BCTechSupport@iridex.com | 2024-08-09 12:11 | Yes | No |  | No |  | No
SRV-RQ001857 | SI | 65600-01 | IQY71339C | CUST-IQ577 PSI, intermittent E06002 over 500mw, request for a PM and provide estimate before repiars | DUPLICATED CUSTOMER COMPLAINT EO6002 HEAD SENT TO H.R.JS |  | BCTechSupport@iridex.com | 2024-08-19 08:51 | Yes | No |  | No |  | No
SRV-RQ001874 | PSD | 70292 | YEZ14136TC | CUST-TxCell, PSD, LED KEEP BLINKING, (IQ577 Included) | GALVO'S AND COMPUTER BOARD ARE NOT WORKING |  | BCTechSupport@iridex.com | 2024-08-22 09:09 | Yes | No |  | No |  | No
SRV-RQ001938 | PSD | 70292 | YEZ13233TC | CUST - TXCELL.PSD. PM | UNABLE TO TESTED COMPUTER BAORD NOT WORKING GALVO'S ARE DOWN REV |  | BCTechSupport@iridex.com | 2024-08-22 09:09 | Yes | No |  | No |  | No
SRV-RQ001775 | PSD | 70300 | GH13228TC | CUST-TxCell,PSD, LED KEPT BLINKING | CONTROL BOX NON-STOP BLINKING, BAN COMPUTER BAORD AND, |  | BCTechSupport@iridex.com | 2024-08-22 09:27 | Yes | No |  | No |  | No
SRV-RQ000503 | PSI | 70300 | GH15707TC | CUST-TxCell. PSI. LED blinking | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING | replaced computer board usb cable tested aligned and calibrated | BCTechSupport@iridex.com | 2024-08-22 09:56 | Yes | No |  | No |  | No
SRV-RQ001928 | SI | 12729 | 28732-LS | CUST-SLx, PSI, Need PM for output power and other functions | LOW POWER HEAD SENT TO HEAD ROOM.JS |  | BCTechSupport@iridex.com | 2024-08-23 07:55 | Yes | No |  | No |  | No
SRV-RQ001947 |  | 12729 | 27423 | CUST-SLx, PSD, LPA- LOW POWER | Low power output. JS |  | BCTechSupport@iridex.com | 2024-08-23 08:11 | Yes | No |  | No |  | No
SRV-RQ001727 | WD | 70292D | YEZ13149TC-D | CUST-TxCell, WD, LED keep blinking just repair | UNABLE TO TESTED CONTROL BOX NON-S\TOP BLINKING |  | BCTechSupport@iridex.com | 2024-08-23 08:15 | Yes | No |  | No |  | No
SRV-RQ001922 | PSD | 30903-H | 1045133 | CUST-LIO.PSD.PM | Cannot test due to Twisted and burnt assembly fiber optic, bent control stem, |  | BCTechSupport@iridex.com | 2024-08-23 12:55 | Yes | No |  | No |  | No
SRV-RQ002028 | PSI | 70292 | YEZ13216TC | CUST-TXCELL.PSI.PM | UNABLE TO TESTED GALVO'S ARE BAD SOME OF THE PARTS NEED UPGRADED |  | BCTechSupport@iridex.com | 2024-08-26 09:49 | Yes | No |  | No |  | No
SRV-RQ001993 | PSD | 65900-H500 | K1228 | CUST-LIO.PSD.PM | Cannot test due to loose housin illumination, burt halogen bulb, twisted |  | BCTechSupport@iridex.com | 2024-08-26 10:41 | Yes | No |  | No |  | No
SRV-RQ001850 | PSD | 31214 | GL16250F30 | CUST-GL. PSD. PM. No issue | Power is not stable and LIO high dac out of spec. . Laser head head require Eval |  | BCTechSupport@iridex.com | 2024-08-26 10:56 | Yes | No |  | No |  | No
SRV-RQ001961 | PSD | 70295 | YH15944TC |  | CONTROL BOX NON-STOP BLINKING DUE TO COMPUTER NOT BOOT UP |  | BCTechSupport@iridex.com | 2024-08-26 13:20 | Yes | No |  | No |  | No
SRV-RQ001974 | PSD | 30903-H500 | K1275 | Cust-PM-PSD-Headsheet light not coming on | Cannot tes du to burnt halogen bulb, twisted assembly fiber optic, missing |  | BCTechSupport@iridex.com | 2024-08-27 15:42 | Yes | No |  | No |  | No
SRV-RQ001984 | PSD | 70292 | YEZ14171TC | CUST-TxCell,PSD, LED blinking non-stop | UNABLE TO TEST CONTROL BOX NON-STOP BLINKING |  | BCTechSupport@iridex.com | 2024-08-28 06:44 | Yes | No |  | No |  | No
SRV-RQ001952 | PSD | 30903-H500 | K1941 | CUST-LIO, PSD, LIO DROPPED AND BROKEN | Cannot test due to delivery mirror loose from flexure mirror assembly, missing |  | BCTechSupport@iridex.com | 2024-08-28 10:26 | Yes | No |  | No |  | No
SRV-RQ001975 | PSD | 70297 | EZZ12907TC | CUST-TxCell, PSD, E05111, PLEASE EVALUATE FOR PROBLEM | CONTROL BOX NON-STOP BLINKING DUE TO COMPUTER NOT BOOTUP. |  | BCTechSupport@iridex.com | 2024-08-28 10:56 | Yes | No |  | No |  | No
SRV-RQ001949 | PSD | 65515-H500 | 1362356 | CUST-LIO,PSD, NO ILLUMINATION, BROKEN KNOB | Cannot test due to Twisted and burnt assembly fiber optic, burnt halogen bulb, |  | BCTechSupport@iridex.com | 2024-08-28 11:58 | Yes | No |  | No |  | No
SRV-RQ001982 | PSD | 30903-H500 | 1341748 | CUST-LIO,PSD, VIEWING ASSAY SERATED, BROKEN | Cannot test due to Viewing assembly separated from T-slot adapter, broken stem |  | BCTechSupport@iridex.com | 2024-08-29 09:30 | Yes | No |  | No |  | No
SRV-RQ001969 | PSD | 30903-H | L2518 | CUST-LIO.PSD. PM | Cannot test due to viewing assembly dislocated from T-slot, bunrt halogen bulb, |  | BCTechSupport@iridex.com | 2024-08-29 15:25 | Yes | No |  | No |  | No
SRV-RQ001887 | PSI | 13030 | 27971-LS | CUST-SLX.PSI. low power and customer told me that sometimes unit shuts down without touching. | As customer complaint low power but found damaged display need to |  | BCTechSupport@iridex.com | 2024-09-03 10:18 | Yes | No |  | No |  | No
SRV-RQ002026 | PSD | 13153-H500 | K1868 | CUST-LIO, PSD, OUT OF ALIGNMENT | Cannot test due to loose housing illumination, viewing assembly loose from |  | BCTechSupport@iridex.com | 2024-09-03 11:05 | Yes | No |  | No |  | No
SRV-RQ001990 | PSD | 30903-H500 | 1362320 | CUST-LIO, NO WHITE LIGHT, LOW POWER | Cannot test due to Twisted and burnt assembly fiber optic, burnt halogen bulb |  | BCTechSupport@iridex.com | 2024-09-03 12:25 | Yes | No |  | No |  | No
SRV-RQ001983 | PSD | 87300 | 1104689791 | CUST-LIO.PSD.LOW POWER , | Cannot test due to twisted and broken assemly fiber optic, battery came without |  | BCTechSupport@iridex.com | 2024-09-03 14:02 | Yes | No |  | No |  | No
SRV-RQ001998 | PSD | 13152-H500 | K1098 | CUST-LIO, PSD,BURNT FIBER | Cannot test due to Twisted and burnt assebly fiber optic, missing eyepiece |  | BCTechSupport@iridex.com | 2024-09-03 14:50 | Yes | No |  | No |  | No
SRV-RQ001614 | CEM | 33005-1 | GLX18673C | CUST-GLx. CEM. Calibration due | Stripped Ferrule . Laser head querire Eval. ct |  | BCTechSupport@iridex.com | 2024-09-04 07:12 | Yes | No |  | No |  | No
SRV-RQ001988 | PSD | 30903-H | L3201 | CUST-LIO.PSD.PM | fractured T-slot adapter, missing eyepiece rubber, extension cable transmite low |  | BCTechSupport@iridex.com | 2024-09-04 08:45 | Yes | No |  | No |  | No
SRV-RQ001339 | PSI | 10798 | 2575F30 | CUST-HEAD-SLx,PSI,LPA | Laser head need 30w diode #31096,upgrade heater board 30791, and adapter 10153 |  | BCTechSupport@iridex.com | 2024-09-04 14:26 | Yes | No |  | No |  | No
SRV-RQ002029 | PSI | 13153-H500 | 1351104 | CUST-LIO, PSD, DAMAGED MIRROR AND KNOBS | Cannot test due to Twisted and burn assembly fiber optic, coating delaminated |  | BCTechSupport@iridex.com | 2024-09-04 15:45 | Yes | No |  | No |  | No
SRV-RQ001864 | PSI | 31255-01 | GH12273EV | CUST-SLA. PSI. PM | PM SLA, Unit has low power readings(REPLACE FIBER JUMPER), Missing dust |  | BCTechSupport@iridex.com | 2024-09-05 11:41 | Yes | No |  | No |  | No
SRV-RQ001945 | PSD | 65500 | IQG81030C | CUST-IQ532,PSD, TOUTCH SCREEN WON'T WORK | Need to replace display LCD and encoder for further investigation. SD |  | BCTechSupport@iridex.com | 2024-09-05 14:56 | Yes | No |  | No |  | No
SRV-RQ002078 | PSD | 65900-H500 | 1362356 | CUST-LIO,PSD, NO ILLUMINATION, BROKEN KNOB | Cannot test due to twisted and burnt assembly fiber optic, burnt halogen bulb, |  | BCTechSupport@iridex.com | 2024-09-05 15:08 | Yes | No |  | No |  | No
SRV-RQ001856 | PSD | 31215 | GL16156C30 | CUST-GL. PSD. Connect fiber & reset | Console keep reset, when firing power . Error 5 and Error 24 .Laser head |  | BCTechSupport@iridex.com | 2024-09-06 16:34 | Yes | No |  | No |  | No
SRV-RQ001856 | PSD | 31215 | GL16156C30 | CUST-GL. PSD. Connect fiber & reset | Console keep reset, when firing power . Error 5 and Error 24 .Laser head |  | BCTechSupport@iridex.com | 2024-09-06 13:33 | Yes | No |  | No |  | No
SRV-RQ001779 | PSI | 30905 | GL14115F | CUST-GL. PSI. Low power output | Error 5 Retest and found low power and high dac. |  | BCTechSupport@iridex.com | 2024-09-09 07:27 | Yes | No |  | No |  | No
SRV-RQ001965 | SI | 12729 | 27429-LS | CUST-LIO, PSD, LOW POWER | Low power output. JS |  | BCTechSupport@iridex.com | 2024-09-09 10:01 | Yes | No |  | No |  | No
SRV-RQ002014 | PSI | 30898-CSO | EZC10784 | CUST-SLA.PSI.PM | Burnt Fiber, Output Mirror Loose, Output Mirror bearing is tight. |  | BCTechSupport@iridex.com | 2024-09-09 13:14 | Yes | No |  | No |  | No
SRV-RQ002005 | PSD | 30903-H500 | K2050 | CUST-LIO.PSD. PM | Twisted assembly fiber optic and missing eyepiece ruber, bad +2 diopter lend on | Replaced +2 diopter lens, dust cap with lnayard, eyepiece rubber, repolished | BCTechSupport@iridex.com | 2024-09-10 14:13 | Yes | No |  | No |  | No
SRV-RQ002012 | PSD | 30903-H500 | 1344082 | CUST-LIO.PSD.PM | Cannot tes due to Burnt assembly fiber optic, broken extention cable, bent |  | BCTechSupport@iridex.com | 2024-09-10 14:35 | Yes | No |  | No |  | No
SRV-RQ002068 | PSD | 65515-H500 | K1702 | CUST-LIO.PSD.CONTROL KNOBS ARE BROKEN | Cannot test due to aimig beam clipping, loose viewing assembly from T-slot |  | BCTechSupport@iridex.com | 2024-09-11 08:46 | Yes | No |  | No |  | No
SRV-RQ001781 | PSD | 32000-1 | TX18596C |  | PM check up .Retest and found Error 5 and Error 23 . Low power |  | BCTechSupport@iridex.com | 2024-09-11 09:04 | Yes | No |  | No |  | No
SRV-RQ002034 | PSD | 65900-H500 | 1356323 | CUST-LIO, PSD, LOW POWER, VERIFY SHIPPING ADRESS BEFORE SHIPPING | Cannot test due to Twisted and burnt assembly fiber optic, Burnt halogen bulb, |  | BCTechSupport@iridex.com | 2024-09-11 12:29 | Yes | No |  | No |  | No
SRV-RQ001813 | WI | 70295 | YH15441TC | CUST-TxCell. WI. E09505 again after just being serviced in July 2023. | UNABLE TO TESTED ERROR E09505 |  | BCTechSupport@iridex.com | 2024-09-11 13:28 | Yes | No |  | No |  | No
SRV-RQ002055 | PSD | 30268-H500 | L1621 | CUST-LIO.PSD. bulb not lighting up | T-slot adapter, bent control stem, fractured housing with window and wil need |  | BCTechSupport@iridex.com | 2024-09-11 14:49 | Yes | No |  | No |  | No
SRV-RQ001770 | WD | 32000-4 | TX19426C | CUST-TX, WD, LOW POWERS , PM | Low power . power shuts off when firing. Laser head require Eval . CT | It's fixed will send to console area today. Yo | BCTechSupport@iridex.com | 2024-09-13 07:42 | Yes | No |  | No |  | No
SRV-RQ001926 | SD | 65500 | IQG81029C | CUST-IQ532, PSD, JUMPY ENCODER AND NO ILLUMUNATION | ILLUMINATION LIGHT WORKS FINE THE PROBLEM WITH ENCODER IS JUMPS |  | BCTechSupport@iridex.com | 2024-09-13 08:09 | Yes | No |  | No |  | No
SRV-RQ002048 | PSD | 66245-1 | IQI92811C | CUST-G6, PSD, Dead battery message |  |  | BCTechSupport@iridex.com | 2024-09-13 09:57 | Yes | No |  | No |  | No
SRV-RQ001986 | PSI | 70295 | YH14838TC | CUST - TxCell. PSI. E05111 Error | UNABBLE TO TESTED ONE OFF THE POWER SUPPY IS NOT WORKING, |  | BCTechSupport@iridex.com | 2024-09-13 10:07 | Yes | No |  | No |  | No
SRV-RQ002049 | PSI | 70295 | YH13591TC | CUST-TxCell, PSI, FILTERPLATE CLAMP ASSY LOOSE |  |  | BCTechSupport@iridex.com | 2024-09-16 11:29 | Yes | No |  | No |  | No
SRV-RQ002018 | PSD | 65622 | RYEZ15936 | CUST-SLA,PSD,SLA MIGHT BE BENT? VERYFY SHIPPING ADDRESS BEFORE SHIPPING | Lever Ball Pin Assy is rotated in the wrong direction, Fiber Jumper has low | Rotated the lever ball pin assy in the correct position, Replaced Fiber Jumper, | BCTechSupport@iridex.com | 2024-09-17 08:40 | Yes | No |  | No |  | No
SRV-RQ001616 | CEM | 33005 | GLX18729C | CUST-GLx. CEM. Calibration due | Stripped fiber connector . Require laser head eval. ct |  | BCTechSupport@iridex.com | 2024-09-18 13:19 | Yes | No |  | No |  | No
SRV-RQ001848 | PSD | 30268-H | L2489 | CUST-LIO. PSD. Green flashback. Please check ESF | Cannot test due to Twisted and broken assembly fiber optic, burnt halogen bulb, |  | BCTechSupport@iridex.com | 2024-09-19 15:02 | Yes | No |  | No |  | No
SRV-RQ002056 | SCD | 30903-H500 | K1472 | CUST-LIO.SCD.PM | Cannot test due to Twisted and broken assembly fiber optic, wrong and burnt |  | BCTechSupport@iridex.com | 2024-09-19 15:17 | Yes | No |  | No |  | No
SRV-RQ001911 | PSI | 32002 | TX18897H | CUST-HEAD-TX, PSI, LPA ERROR6 | Need assy vanadate 30264. Yo |  | BCTechSupport@iridex.com | 2024-09-20 07:47 | Yes | No |  | No |  | No
SRV-RQ002008 | PSD | 13153-H | 1300849 | Cust-LIO-PM-PSD | Cannot test due to Twisted and broken assembly fiber optic, broken control stem |  | BCTechSupport@iridex.com | 2024-09-20 08:25 | Yes | No |  | No |  | No
SRV-RQ001943 | PSD | 32000-4 | TX18882C | CUST-TX, PSD, ERROR 38 | Error 5 and error 38 .Low power and high dac .Lasrer head require Eval. CT | It's fixed, realign laser head to manufacture TX specifiction. Will send to | BCTechSupport@iridex.com | 2024-09-20 09:22 | Yes | No |  | No |  | No
SRV-RQ002080 | PSD | 30903-H | 1334102 | CUST-LIO, PSD, NEED SERVICE ONLY PROBLEM , ADRESS VERIFIED | Cannot test due to Twisted and burnt assembly fiber optic, broken control stem, |  | BCTechSupport@iridex.com | 2024-09-20 09:33 | Yes | No |  | No |  | No
SRV-RQ001939 | PSD | 12093 | RH7859-LS | CUST. SLA.PSD.PM | No problems found. PM, CLean mirrors,lenses and test. | PM, Cleaned Mirrors, Lenses and Tested. | BCTechSupport@iridex.com | 2024-09-23 09:23 | Yes | No |  | No |  | No
SRV-RQ002083 | PSD | 31214 | GL14439F30 | CUST-GL, PSD, Heating/colling, Verify Ship To, provide estimate before repairs | Heating Delta sctuck can't firing power . Error 34 and error 21 .Smell smoke ins |  | BCTechSupport@iridex.com | 2024-09-23 12:17 | Yes | No |  | No |  | No
SRV-RQ002030 | PSD | 31172-01 | GH16246LM | CUST-SLA, PSD, BROKEN FIBER AND NEED ALIGNMENT | Broken Fiber Jumper. Clamp knob not attached to pivot base, |  | BCTechSupport@iridex.com | 2024-09-24 11:09 | Yes | No |  | No |  | No
SRV-RQ002093 | PSD | 70292 | YEZ14168TC | Cust-TxCell, PSD, No details of the error No provided, please investigate ( an IQ577 is included) | CANNOT DUPLICATED THE PROBLEM. | TESTED CLEAN AND INSPECTED. | BCTechSupport@iridex.com | 2024-09-25 09:41 | Yes | No |  | No |  | No
SRV-RQ002072 | PSD | 70292D | YEZ13254TC-D | CUST-TxCell, PSD, TxCell does not boot ( non-stable blinking LED) | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING |  | BCTechSupport@iridex.com | 2024-09-25 12:46 | Yes | No |  | No |  | No
SRV-RQ001963 | SI | 13030-220 | 28909-LS | CUST-SLx, PSI, PM and check for footpedal stuck per user | UNIT CAME FOR PM AND FOOT SWITCH FOOT SWITCH CONNECTOR WAS |  | BCTechSupport@iridex.com | 2024-09-26 09:36 | Yes | No |  | No |  | No
SRV-RQ002144 | PSD | 65900-H500 | 1355855 | CST-LIO.PSD. KNOB LOOSE IN BACK OF LIO | Cannot test due to twisted and burnt assembly fiber optic, burnt Heine halogen |  | BCTechSupport@iridex.com | 2024-09-30 12:17 | Yes | No |  | No |  | No
SRV-RQ001933 | PSD | 65515-H500 | 1360538 | CUST-LIO, PSD, LOW POWER | Cannot test due to Twisted and broken assembly fiber optic, bent control stem, |  | BCTechSupport@iridex.com | 2024-09-30 12:44 | Yes | No |  | No |  | No
SRV-RQ002011 | PSD | 70292 | YEZ14616TC | CUST-TxCell.PSD. PM | NO PROBLEM FOUND. |  | BCTechSupport@iridex.com | 2024-09-30 13:23 | Yes | No |  | No |  | No
SRV-RQ002130 | PSD | 30903-H500 | K2141 | CUST-LIO, PSD, UNABLE TO FOCUS DIFFRENT IN BOTH OCCULAR | Cannot testd due to scratch and chipped input assemble fiber optic, burnt Heine |  | BCTechSupport@iridex.com | 2024-09-30 14:58 | Yes | No |  | No |  | No
SRV-RQ002179 | PSI | 70292 | YEZ13030TC | CUST-TXCELL.PSI. PM NEED ESTIMATE BEFORE REPAIR | GALVO AND COMPUTER BOARD ARE BAD, MIRROR LASER IS MISSING |  | BCTechSupport@iridex.com | 2024-10-02 08:13 | Yes | No |  | No |  | No
SRV-RQ002096 | PSI | 70292 | YEZ15822TC | CUST-TxCell, PSI, MS framework.NET crashed. Provide estimate before completion | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING |  | BCTechSupport@iridex.com | 2024-10-02 11:01 | Yes | No |  | No |  | No
SRV-RQ002178 | PSI | 70292 | YEZ13031TC | CUST-TXCELL.PSI, PM. PLEASE PROVIDE ESTIMATE BEFORE REPAIR | CONTROL BOX NON-STOP BLINKING,COMPUTER BOARD NOT WORKING |  | BCTechSupport@iridex.com | 2024-10-02 12:45 | Yes | No |  | No |  | No
SRV-RQ002170 | PSD | 65515-H | 1344070 | Cust-LIO, PSD, Broken handel, Please deliver to AZ after recieving ( HOT, URGENT) | fractured, burnt halogen bulb, mising control stem, eyepiece rubber fractured |  | BCTechSupport@iridex.com | 2024-10-02 15:24 | Yes | No |  | No |  | No
SRV-RQ001746 | PSD | 65777 | RC0344 | CUST-REMOTE. PSD. Returned w/ RMA#1096 | The remote came back with RMA1096. JS |  | BCTechSupport@iridex.com | 2024-10-03 06:48 | Yes | No |  | No |  | No
SRV-RQ002182 | PSI | 70292 | YEZ12908TC | CUST-TXCELL.PSI. PM NEED ESTIMATE BEFORE REPAIR | CONTROL BOX NON-STOP BLINKING, COMPUTER BOARD NOT WORKING |  | BCTechSupport@iridex.com | 2024-10-03 07:55 | Yes | No |  | No |  | No
SRV-RQ002044 | SI | 13030-220 | 29808-LS | CUST-SLx, PSI, "NO SHOTS RECOCNIZED JUST REPLACED POWER SUPPLY | NOT SHOT RECOGNISE I HAVE NO CLUE WHAT THAT MEAN BUT FOUND | Reolaced fiber adapter and realign laser head. | BCTechSupport@iridex.com | 2024-10-04 08:28 | Yes | No |  | No |  | No
SRV-RQ002059 | PSI | 70295 | YH15286TC | CUST-TXCELL.PSI.POWER SUPPLY IS DAMAGE (12V) | POWER SUPPLU IS DAMAGE 12V |  | BCTechSupport@iridex.com | 2024-10-07 09:34 | Yes | No |  | No |  | No
SRV-RQ001906 | PSI | 65600-01 | IQY71484C | CUST-IQ577, PSI E06004 (CAME WITH TXCELL CONSOLE) | Need to replace encoder. JS | Received encoder. DN | BCTechSupport@iridex.com | 2024-10-10 09:05 | Yes | No |  | No |  | No
SRV-RQ002164 | PSD | 30903-H | 1143518 | CUST-LIO, PSD, DROPPED AND IMAGE NOT CENTERED | Cannot test due to Twisted and burt halogen bulb, loose cap fiber adapter from |  | BCTechSupport@iridex.com | 2024-10-10 13:01 | Yes | No |  | No |  | No
SRV-RQ002207 | SCD | 65515-H500 | 1362985 | CUST-LIO, CHECK FLASHBACK ISSUE AND DATA SHEET AVAILABLE | Cannot test due to broken control stem, loose viewing assembly from T-slot |  | BCTechSupport@iridex.com | 2024-10-10 14:44 | Yes | No |  | No |  | No
SRV-RQ002109 | SI | 65600-01 | IQY71132C | CUST-IQ577, PSI, Heating delta displays randomly. Need an estimate | E03020 @START UP HEAD SENT TO HEAD ROOM.JS |  | BCTechSupport@iridex.com | 2024-10-11 10:00 | Yes | No |  | No |  | No
SRV-RQ002215 | PSD | 30903-H | 1171620 | CUST-LIO,PSD,BROKEN KNOBS | Cannot test due to Twisted and broken assembly fiber optic, burnt halogen bulb, |  | BCTechSupport@iridex.com | 2024-10-14 10:44 | Yes | No |  | No |  | No
SRV-RQ002157 | SD | 13030 | 30134-LS | CUST-SLx, PSD, NO ILLUMINATION LIGHT | NO ILLUMINATION NEED POWER SUPPLY AND RECALIBRATEUNIT.JS |  | BCTechSupport@iridex.com | 2024-10-15 12:50 | Yes | No |  | No |  | No
SRV-RQ002194 | PSD | 70297 | EZZ13063TC | Cust-IQ532/TxCell, PSD, No communicatin between the IQ& TxCell E05111 | UNABLE TOTESTED CONTROL BOX NON-STOP BLINKING DO TO COMPUTER BOARD NOT WORKING |  | BCTechSupport@iridex.com | 2024-10-18 06:53 | Yes | No |  | No |  | No
SRV-RQ002236 | PSD | 30903-M | L3499 | CUST-LIO, PSD, UNABLE TO FIND SN. SHIPPING LIO | Cannot test due to Twisted and broken assembky fiber optic, wrong control stem, |  | BCTechSupport@iridex.com | 2024-10-21 10:57 | Yes | No |  | No |  | No
SRV-RQ002198 | PSD | 30903-H | 1305737 | CUST-LIO, PSD, LOW POWER, BROKEN | Cannot test due to Damaged SMA connector, missing PC contact, wrong halogen bulb |  | BCTechSupport@iridex.com | 2024-10-21 11:57 | Yes | No |  | No |  | No
SRV-RQ002089 | PSI | 65600-01 | IQY71286C | CUST-IQ577, PSI, E06010 "Laser power output out of range" E00000 "Failure | Power output out of range. SD |  | BCTechSupport@iridex.com | 2024-10-21 13:53 | Yes | No |  | No |  | No
SRV-RQ002002 | PSD | 30658 | GL14228C30 | CUST-GL, PSD, LPA | Rotation has problem and powers are not stable cause low power . Laser head | Realign laser head to manufacture GL 3rd cooler specification. Yo | BCTechSupport@iridex.com | 2024-10-21 18:10 | Yes | No |  | No |  | No
SRV-RQ002146 | SD | 30268-M | L2621 | CUST-LIO, PSD, LOW PWER, DAMADGED FIBER | Cannote test due to Burnt assembly fiber optic, broken control stem, viewing |  | BCTechSupport@iridex.com | 2024-10-22 08:02 | Yes | No |  | No |  | No
SRV-RQ001715 | PSD | 30905 | GL13766F | CUST-GL. PSD. Stripped ferrule | SRTIPPED FERRULER , LASER CAN'T RECOGNIZE THE FIBER.LASER HEAD REQUIRE EVAL. CT | It's fixed, will move to console. | BCTechSupport@iridex.com | 2024-10-22 14:20 | Yes | No |  | No |  | No
SRV-RQ002154 | SD | 32000-4 | TX19527C | CUST-TX, WD, JUST REPAIRED FOR nOT RECOGNIZING FIBER PLEASE EVALUATE | INTERMITTENGLY NOT RECOGNISE FIBER.JS | MADE THIS CONTACT FIRM AND RETESTED. JS | BCTechSupport@iridex.com | 2024-10-23 10:09 | Yes | No |  | No |  | No
SRV-RQ002274 | PSD | 32000D-4 | TX18882C-D | CUST-TX, PSD, ERROR 38 | Error 5 and error 38 .Low power and high dac .Lasrer head require Eval. CT | Realigned laser head to manufacture TX spec. YO | BCTechSupport@iridex.com | 2024-10-23 12:06 | Yes | No |  | No |  | No
SRV-RQ001916 | PSI | 30904 | 2575F30 | CUST-HEAD-SLx,PSI,LPA | Laser head need 30w diode #31096,upgrade heater board 30791, and adapter 10153 |  | BCTechSupport@iridex.com | 2024-10-23 13:50 | Yes | No |  | No |  | No
SRV-RQ002117 | PSD | 66245D-1 | IQI92002C-D | Cust-G6-PM-PSD-Needs Estimate Prior to Work | PCBA encoder and module A/C input are broken.SD | Replace PCBA encoder and module input A/C | BCTechSupport@iridex.com | 2024-10-24 15:56 | Yes | No |  | No |  | No
SRV-RQ002038 | PSI | 32000-4 | TX18839C | CUST-TX.PSD. Not providing enough Power output | Error 5 and error 40 . power shuts off when firing ar hight power 2500mw. | Realign laser head to manufacture TX specification. Yo | BCTechSupport@iridex.com | 2024-10-25 06:49 | Yes | No |  | No |  | No
SRV-RQ002194 | PSD | 65500 | IQG81269C | Cust-IQ532/TxCell, PSD, No communicatin between the IQ& TxCell E05111 | Power output low .SD |  | BCTechSupport@iridex.com | 2024-10-25 11:02 | Yes | No |  | No |  | No
SRV-RQ002147 | PSD | 30685 | GH7267M | CUST-SLA, PSD, LOW PWER, DAMADGED FIBER | BURNT FIBER CAUSING LOW POWER, ESF FILTERS IS SCRATCHED, MISSING |  | BCTechSupport@iridex.com | 2024-10-25 13:05 | Yes | No |  | No |  | No
SRV-RQ002137 | PSD | 13030 | 30350-LS | CUST-SLx, PSD, PM | Customer return for PM check. SD | Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2024-10-25 15:16 | Yes | No |  | No |  | No
SRV-RQ002040 | PSD | 32000-1 | TX1329REVB | CUST-TX, PSD, EMERG. STOP AND ERROR 2 | Error 2 P/S board bad . photocell 1 and 2 are not matching ,AND ENT hight dac | Realign laser head to manufacture TX specification. Yo | BCTechSupport@iridex.com | 2024-10-26 09:28 | Yes | No |  | No |  | No
SRV-RQ002213 | PSD | 13030 | 30198-LS | CUST-SLX.PSD.PM | Cust-return for PM check. SD | Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2024-10-26 09:27 | Yes | No |  | No |  | No
SRV-RQ002228 | PSD | 32000-1 | TX1184REVB | CUST-TX, PSD, ERROR 2 POWER SUPP ERROR | ERROR 2 IN MEMORY CHECK POWER SUPPLY VOLTAGE WAS DISPLAYED AND | READJUSTED TEMPRATURE TO REDUCE NOISE AND RECALIBRATED | BCTechSupport@iridex.com | 2024-10-26 09:37 | Yes | No |  | No |  | No
SRV-RQ002010 | PSD | 65600-01 | IQY71897C | CUST-IQ577.PSD. aiming beam is dimmer than Port 2 | aiming beam is dimmer than Port 2. Port 1 no power. SD | -Replaced fine adjust screws, realigned the laser head, recalibrated the unit. | BCTechSupport@iridex.com | 2024-10-26 09:43 | Yes | No |  | No |  | No
SRV-RQ002023 | PSI | 70295 | YH15099TC | CUST-TXCELL.PSI. AIMING LIGHT IS CRACKED | BURNED FIBER,BROKEN FLEX PLATE THAT'S WHY THE AIMING BEAM IS NOT ON THE CENTER. | REPLACED INTERFACE BAORD USB CABLE FLEX PLATE, POLISH THE FIBER | BCTechSupport@iridex.com | 2024-10-26 09:53 | Yes | No |  | No |  | No
SRV-RQ002126 | PSD | 65600D-01 | IQY71205C-D | Cust-IQ577, PSD, No details of the error No provided, please investigate ( an TxCell is included) |  |  | BCTechSupport@iridex.com | 2024-10-26 10:03 | Yes | No |  | No |  | No
SRV-RQ002120 | PSD | 32000-1 | TX1165REVB | CUST-TX,PSD,PM CHECKUP, ADRESS VERIFIED | PM CHECK | RETESTED ALL THE DELIVERY DEVICES POWER WITH IN SPEC.JS | BCTechSupport@iridex.com | 2024-10-26 10:33 | Yes | No |  | No |  | No
SRV-RQ001980 | PSD | 32000-1 | TX1084REVB-D | CUST-TX,PSD, HEATING DELTA INTERMITTEN TLY | heating delta , head temperture rurn up and down very very loud .Clicking noise | Realign laser head to manufacture TX specification. Yo | BCTechSupport@iridex.com | 2024-10-26 10:34 | Yes | No |  | No |  | No
SRV-RQ001784 | PSD | 66245-1 | IQI92537C | CUST-G6. PSD. PM. No issue | PM CHECK, POWER WITH IN SPEC. | RETESTED ALL THE DELIVERY DEVICES JS | BCTechSupport@iridex.com | 2024-10-26 10:48 | Yes | No |  | No |  | No
SRV-RQ001981 | PSD | 65500 | IQG81027C | CUST-IQ532, PSD, JUMPY ENCODER | jumpy encoder as customer complaint interval knob jumps.js | REPLACED ENCODER AND UPGRADED FFC CABLE AND RETESTED.JS | BCTechSupport@iridex.com | 2024-10-26 10:49 | Yes | No |  | No |  | No
SRV-RQ002145 | WD | 13030 | 30297-LS | CUST-SLx, WD, NO BEEPER OR SOUND | Fan cable install wrong location J53. sd | Reinstall fan cable in to J9 and retested all the delivery devices. SD | BCTechSupport@iridex.com | 2024-10-28 07:52 | Yes | No |  | No |  | No
SRV-RQ002142 | WD | 66120 | IQHG81047C | CUST-IQHG-532,WD, E05111 AND LOW POWER | No problem found. SD | Retested all the delivery devices, SD | BCTechSupport@iridex.com | 2024-10-28 08:50 | Yes | No |  | No |  | No
SRV-RQ002247 | PSD | 31172-02 | GH16090LM | CUST-SLA,PSD, LOW AIMBEAM AND IMAGE NOT CENTERED | Fiber Jumper was not attached to SLA, ESF Filter is Scratched, Align and |  | BCTechSupport@iridex.com | 2024-10-28 09:24 | Yes | No |  | No |  | No
SRV-RQ001833 | PSI | 66245-1 | IQI93295C | CUST-G6. PSI. Low power output by 35% | Low power output .SD | Realign and calibration | BCTechSupport@iridex.com | 2024-10-29 05:20 | Yes | No |  | No |  | No
SRV-RQ002053 | PSD | 31215D | GL15138C30-D | CUST-GLx, PSD, MULTIPLE ERRORS PER CUSTOMER NO OTHER PROBLEM GIVEN | Error 5 and Error 23 low power . Laser head need Eval ct |  | BCTechSupport@iridex.com | 2024-10-29 06:21 | Yes | No |  | No |  | No
SRV-RQ002261 | PSD | 65515-H500 | 1354782 | CUST-LIO, PSD, NO WHITE LIGHT, BULB REPLACED STIIL NO LIGHT |  | Cannot test due to Twisted and burnt assebly fiber optic, wrong and burnt | BCTechSupport@iridex.com | 2024-10-29 12:45 | Yes | No |  | No |  | No
SRV-RQ002104 | PSD | 65500D | IQG81616C-D | CUST-IQ532,PSD,INTERMITTENLY DROPS TREAT TO STD BY | port #1 not detect , power output low. | Replace PCBA resistance fiber detect. SD | BCTechSupport@iridex.com | 2024-10-30 07:13 | Yes | No |  | No |  | No
SRV-RQ001989 | PSD | 32000-1 | TX18142C | CUST-TX, PSD, Error 2, power supply error | CUSTOMER COMPLAINT ERROR 2 HAVE NOT DUPLICATED ERROR.JS | READJUSTED TEMPRATURE AND RECALIBRATED DELIVERY DEVICES.JS | BCTechSupport@iridex.com | 2024-10-30 08:01 | Yes | No |  | No |  | No
SRV-RQ002039 | PSI | 65500-08 | IQG82018C | CUST-IQ532.PSI. Port Switching error | Power output out of range. SD | -Realigned the laser head, recalibrated the unit. Verfied all the function , | BCTechSupport@iridex.com | 2024-10-31 06:28 | Yes | No |  | No |  | No
SRV-RQ002153 | SD | 13030 | 26041-LS | CUST-SLX.PSD.Broken emergency button and unit display blinking on and off | DISPLAY DAMAGED AFTER REPAIR DISPLAY FOUND LOW POWER, |  | BCTechSupport@iridex.com | 2024-10-31 10:24 | Yes | No |  | No |  | No
SRV-RQ001917 | SD | 32000-4 | TX19455C | CUST-TX, WD, HEETING DELTA | Heating delta. Error 5 low power, keep resetting. CT | Realign laser head to manufacture TX specification.Yo | BCTechSupport@iridex.com | 2024-10-31 14:13 | Yes | No |  | No |  | No
SRV-RQ002216 | SCD | 65515-H500 | 1362432 | CUST-LIO.SCD.PM | Cannot test due to broken control stem, burnt halogen bulb, fractured T-slot |  | BCTechSupport@iridex.com | 2024-11-05 11:06 | Yes | No |  | No |  | No
SRV-RQ002222 | PSD | 87300 | 1104007138 | CUST-LIO, PSD, NO AIM BEAM AND NEED EVALUATION, VERIFY SHIPPING ADRESS BEFORE SHIPPING | Cannot test due to broken assebly fiber optic, will need adjust and alignment. |  | BCTechSupport@iridex.com | 2024-11-05 11:22 | Yes | No |  | No |  | No
SRV-RQ002256 | PSD | 30903-H500 | 1351113 | CUST-LIO.PSD BROKEN KNOB MIRROR | Cannot test due to Twisted and burnt assembly fiber optic, burnt halogen bulb, |  | BCTechSupport@iridex.com | 2024-11-05 12:21 | Yes | No |  | No |  | No
SRV-RQ002226 | PSI | 87300 | 1100906951 | CUST-LIO.PSI. PM | Cannot test due to Twisted and broken assembly fiber optic will need adjusment |  | BCTechSupport@iridex.com | 2024-11-05 14:17 | Yes | No |  | No |  | No
SRV-RQ002293 | PSD | 87300 | 1101319504 | CUST-LIO,PSD, DAMEGED FIBER OPTIC AND POWER SUPPLY CABLE | Cannot test due to twisted and broken assebly fiber optic, damaged LIO charger |  | BCTechSupport@iridex.com | 2024-11-05 15:42 | Yes | No |  | No |  | No
SRV-RQ001956 | PSD | 65600-08 | IQY72483C | Cust-IQ577, PSD, E05111, PM, Eval. needed (TxCell returning with the console)) | LOOKS LIKE CONSOLE DROPED DAMAGED CHASSIE. UNABLE TO CONNECT FIBER | REOPEN FRONT PANEL REALIGNED AND REPLACEDRFID INTERFACE ASSY | BCTechSupport@iridex.com | 2024-11-05 17:16 | Yes | No |  | No |  | No
SRV-RQ002094 | SD | 65500 | IQG81193C | CUST-IQ532,PSD, JUMPY ENCODER | CUSTOMER COMPLAINT BAD ENCODER,DUPLICATED AND BACK PANEL |  | BCTechSupport@iridex.com | 2024-11-06 07:17 | Yes | No |  | No |  | No
SRV-RQ002312 | PSI | 13153-H500 | K1993 | CUST-LIO.PSD.PM | Cannot test due to Twisted and burnt assembly fiber optic, loose housing |  | BCTechSupport@iridex.com | 2024-11-06 08:39 | Yes | No |  | No |  | No
SRV-RQ002316 | PSI | 65600-02 | IQY71132C | CUST-IQ577, PSI, Heating delta displays randomly. Need an estimate | E0320 during star up. Laser head need evaluation. JS |  | BCTechSupport@iridex.com | 2024-11-06 12:59 | Yes | No |  | No |  | No
SRV-RQ002313 | PSD | 13152-H | 1170219 | CUST-LIO.PSD.PM HOT! | Cannot test due to missing assembly fiber optic, extension cable, halogen bulb, |  | BCTechSupport@iridex.com | 2024-11-06 13:13 | Yes | No |  | No |  | No
SRV-RQ002234 | PSD | 31214D | GL15577F30-D | CUST-GLx, PSD, HEATING DELTA JUST REPAIR FOR THE SAME PROBLEM | Need to replace a power supply, 31202 but the part is obsoleted. CT |  | BCTechSupport@iridex.com | 2024-11-06 13:22 | Yes | No |  | No |  | No
SRV-RQ002227 | PSD | 30903-H500 | 1380000170 | CUST-LIO.PSD.PM | Cannot test due to Twisted and broken assembly fiber optic, damaged extension |  | BCTechSupport@iridex.com | 2024-11-07 07:48 | Yes | No |  | No |  | No
SRV-RQ001973 | SCD | 66245-1 | IQI93216C | Cust-G6-PM-SCD-PP | PM CHECK, FOOT SWITCH CAME WITH UNIT.JS | RETESTED ALL THE DELIVERY DEVICES, AND FOOT SWITCH TESTED AS WELL.JS | BCTechSupport@iridex.com | 2024-11-07 09:42 | Yes | No |  | No |  | No
SRV-RQ002172 | PSI | 13030-220 | 29882-LS | CUST-SLx , PSI, output is low by 10% or lower | Low power output. |  | BCTechSupport@iridex.com | 2024-11-08 10:15 | Yes | No |  | No |  | No
SRV-RQ002197 | SD | 30905 | GL14455F | Cust-GL.PSD.Pm | UNIT CAME FOR PM FOUND LOW POWER ACTUAL,JS |  | BCTechSupport@iridex.com | 2024-11-11 14:11 | Yes | No |  | No |  | No
SRV-RQ002307 | SCD | 65900-H500 | K1618 | CUST-LIO,SCD, lOOSE INSIDE (LENS) | Cannot test dur to Loose delivery mirror from flexure mirror support, loose |  | BCTechSupport@iridex.com | 2024-11-12 09:18 | Yes | No |  | No |  | No
SRV-RQ002200 | PSD | 31214 | GL15714F30 | CUST-GL, PSD ( SV101408, PM $2800 flat), provide estimate | PM CHECK, NEED RECALIBRATION.JS | READJUSTED TEMPRATURE AND RECALIBRATED ALL THE DELIVERY DEVICES | BCTechSupport@iridex.com | 2024-11-12 14:39 | Yes | No |  | No |  | No
SRV-RQ002206 | PSD | 13152-H | L2844 | CUST-LIO, PSD, No AIM BEAM MAY BROKEN FIBER | Connot test dut to Twisted and broken assembly fiber optic, burnt halogen bulb, | Replaced assembly fiber optic, halogen bulb, dust cap with lanyard, Heine | BCTechSupport@iridex.com | 2024-11-12 15:12 | Yes | No |  | No |  | No
SRV-RQ002336 | SD | 13030 | 27766-LS | CUST-SLX.PSD.PM (WAS BROUGHT IN ADVANTE UNDER 1113262, CUSTOMER FINALLY RESPOND AND NEED TO COMSUME) |  |  | BCTechSupport@iridex.com | 2024-11-14 07:50 | Yes | No |  | No |  | No
SRV-RQ002225 | PSD | 31172-02 | GH11475LM | CUST-SLA,PSD,LOW PWER, VRIFY SHIPPING ADRESS BEFORE SHIPPING | Burned fiber and multiple parts are damaged. | Replaced fiber, and parts. Test ,clean and inspec. | BCTechSupport@iridex.com | 2024-11-14 11:17 | Yes | No |  | No |  | No
SRV-RQ002236 | PSD | 30903-H | L3499 | CUST-LIO, PSD, UNABLE TO FIND SN. SHIPPING LIO, BROKEN KNOBS, LOW POWER | Cannot test duet to Twisted and broken assembly fiber optic, wrong control stem, | Replaced assembly fiber optic, control stem, eccentric shaft, Heine handband, | BCTechSupport@iridex.com | 2024-11-14 13:42 | Yes | No |  | No |  | No
SRV-RQ002134 | SI | 32000-1 | TX18104C | CUST TX.PSI.POWER DROP MORE THAN 30% | LOW POWER AS CUSTOMER COMPLAINT HEAD SENT TO HEAD ROOM.JS |  | BCTechSupport@iridex.com | 2024-11-15 06:57 | Yes | No |  | No |  | No
SRV-RQ002243 | PSD | 65500-08 | IQG81853C | CUST-IQ532, PSD,NOT ACCEPTING FIBER ON BOTH PORT | CUSTOMER COMPLAINT FIBER NOT RECOGNISE HAVE NOT DUPLICATED.JS | RETESTED ALL THE DELIVERY DEVICES WITHOUT ANY PROBLEM.JS | BCTechSupport@iridex.com | 2024-11-15 07:46 | Yes | No |  | No |  | No
SRV-RQ002184 | PSI | 70292 | YEZ15668TC | CUST-IQ577, PSI ,PER PHYSICIAN LOW POWER | FIBER HAS LOW POWER. | REPLACED FIBER TESTED ALIGNED CALIBRATED CLEAN AND INSPECTED. | BCTechSupport@iridex.com | 2024-11-15 08:12 | Yes | No |  | No |  | No
SRV-RQ002199 | WD | 65600-08 | IQY72753C | CUST-IQ577,WD, HEARING WEIRD WHEN FOOTSWITCH PRESSED |  |  | BCTechSupport@iridex.com | 2024-11-18 14:18 | Yes | No |  | No |  | No
SRV-RQ002233 | PSD | 30903-H | 1264706 | CUST-LIO.PSD.PM | Cannot test due to Twisted and broken assembly fiber optic, burnt halogen bulb | Replaced assembly fber optic, control stem, eccentric shaft, halogen bulb, dust | BCTechSupport@iridex.com | 2024-11-18 15:57 | Yes | No |  | No |  | No
SRV-RQ002271 | SCD | 66245 | IQI91300C | CUST-SCD.G6. PM | UNIT CAME FOR PM,FOUND POWER OUT OF RANGE,SENT TO | Realign and calibration . | BCTechSupport@iridex.com | 2024-11-19 08:41 | Yes | No |  | No |  | No
SRV-RQ002102 | STG | 66245 | IQI93693C | Both G6 are from Iridex Germany and have low output by 15% and 20% | Power low but still in range . Intermittent dead battery apear . CT | Replaced battery #66331 .Recalirated all power are good with in MFG | BCTechSupport@iridex.com | 2024-11-19 12:14 | Yes | No |  | No |  | No
SRV-RQ002272 | SCD | 70297 | EZZ13142TC | CUST-TxCell,SCD,PLEASE CHECK FLASH BACK AND PROVIDE READING TO CUST |  |  | BCTechSupport@iridex.com | 2024-11-19 12:50 | Yes | No |  | No |  | No
SRV-RQ002339 | PSD | 13030 | 29719-LS | CUST-SLx, PSD, FOOTSWITCH STUCK MESSAGE CHECKED W/ TWO FS | CUSTOMER COMPLAINT FOOT SWITCH STUCK MESSAGE | RECALIBRATED ALL THE DELIVERY DEVICES AND FOOT SWITCH TESTED | BCTechSupport@iridex.com | 2024-11-19 13:13 | Yes | No |  | No |  | No
SRV-RQ002268 | PSD | 30903-H | 1319881 | CUST-LIO.PSD.LIGHT IS NOT FOCUSING | Cannot test due to dislocated flexure mirror loose cap with fiber adapter from |  | BCTechSupport@iridex.com | 2024-11-20 08:34 | Yes | No |  | No |  | No
SRV-RQ001832 | PSI | 66245-1 | IQI91957C | CUST-G6. PSI. UIM application failed to load | DAMAGED ENCODER NEED TO REPLACE BEFORE START TEST | REPLACED ENCODER AND RETESTED ALL THE DELIVERY DEVICES.JS | BCTechSupport@iridex.com | 2024-11-20 11:05 | Yes | No |  | No |  | No
SRV-RQ001962 | WD | 30903-H500 | K2344 | CUST-LIO, WD, BROKEN ILLUMINATION CABLE | Cannot test due to missing assembly fiber optic, extension cable, Heine halogen |  | BCTechSupport@iridex.com | 2024-11-20 11:16 | Yes | No |  | No |  | No
SRV-RQ002241 | PSI | 70295D | YH14212TC-D | CUST-TxCell, PSI, E09501 | control box non-stop blinking, computer board not working. | replaced computer baord usb cable tested calibrated clean and inspected. | BCTechSupport@iridex.com | 2024-11-21 10:04 | Yes | No |  | No |  | No
SRV-RQ002242 | PSI | 70295 | YH14298TC | CUST-TxCell, PSI, NOT CONNECTING TO IQ577 | CANNOT DUPLICATED THE PROBLEM. | TESTED CALIBRATED CLEAN AND INSPECTED. | BCTechSupport@iridex.com | 2024-11-21 10:43 | Yes | No |  | No |  | No
SRV-RQ002047 | SI | 32000-1 | TX18482C | CUST-TX, PSI, NO DISPLAY, BAD DISPLAY PCB TO REPLACED | Not boot up, bad display per customer reported. JS | Replaced controller and display parts. | BCTechSupport@iridex.com | 2024-11-21 11:04 | Yes | No |  | No |  | No
SRV-RQ002343 | CEM | 65600-01 | IQY71612C | IQ577.CEM.DAME PROJECT FOR JOAN | DAME project per Joan requested. JS | -Replaced fine adjust screws, realigned the laser head, recalibrated the unit. | BCTechSupport@iridex.com | 2024-11-21 12:10 | Yes | No |  | No |  | No
SRV-RQ001846 | WD | 66245 | IQI91323C | CUST-G6. PSD. patient burn w/ pn:15522-1 lot#801158 | patient burn w/ pn:15522-1 lot#801158. SD | No problem found for this console during MFG tested process.SD | BCTechSupport@iridex.com | 2024-11-22 06:47 | Yes | No |  | No |  | No
SRV-RQ002375 | WD | 65600D-08 | IQY81757C-D | DAME PROJECT FOR JOAN - IQ577 | DAME PROJECT FOR JOAN. Return for PM check.SD | Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2024-11-22 06:52 | Yes | No |  | No |  | No
SRV-RQ002314 |  | 32000-4 | TX19534C | CUST-TX, PSD, PM, Laser under warranty but PM is chargable | PM CHECK, FOUND LOW POWE HEAD SENT TO HEAD ROOM FOR. JS | Realign laser head to manufacture TX specification. Yo | BCTechSupport@iridex.com | 2024-11-22 07:09 | Yes | No |  | No |  | No
SRV-RQ002373 | WD | 65600D-08 | IQY81758C-D | DAME PROJECT FOR JOAN - IQ577 | DEMO RETURN.JS | RETESTED ALL THE DELIVERY DEVICES JS | BCTechSupport@iridex.com | 2024-11-22 07:21 | Yes | No |  | No |  | No
SRV-RQ002355 | SD | 13030-L | 29553-LS-L | LOANER-SLX.LD.BRING BACK INTO INVENTORY (LOANER WAS UNDER RMA01946) |  |  | BCTechSupport@iridex.com | 2024-11-22 08:05 | Yes | No |  | No |  | No
SRV-RQ002282 | WD | 13030 | 30340-LS | CUST-SLx, WD, NO ILLUMINATION | NO ILLUMINATION LIGHT | REPLACED POWER SUPPLY AND RECALIBRATED.JS | BCTechSupport@iridex.com | 2024-11-22 10:12 | Yes | No |  | No |  | No
SRV-RQ000244 | PSD | 66245-1 | IQI92775C | CUST-G6. PM. No problem reported | Broken encoder. SD | Replace encoder and retested all the delivery devices. SD | BCTechSupport@iridex.com | 2024-11-26 09:06 | Yes | No |  | No |  | No
SRV-RQ002301 | PSD | 66245 | IQI91627C | CUST-CYCLO-G6, PSD, PM CHECKUP | Customer return for PM check. SD | Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2024-11-26 14:15 | Yes | No |  | No |  | No
SRV-RQ002334 | PSD | 31172-01 | GH10330LM | CUST.SLA.PSD. Fiber Cable broken at end |  |  | BCTechSupport@iridex.com | 2024-12-02 12:53 | Yes | No |  | No |  | No
SRV-RQ002347 | PSD | 65511 | GH13669RM | CUST-SLA.PSD.PM |  |  | BCTechSupport@iridex.com | 2024-12-03 09:56 | Yes | No |  | No |  | No
SRV-RQ002186 | PSD | 70292 | YEZ13306TC | CUST-TxCell.SCD.faulty TxCell and the IQ console (TxCell) | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING COMPUTER BOARD NOT WORKING |  | BCTechSupport@iridex.com | 2024-12-03 09:56 | Yes | No |  | No |  | No
SRV-RQ002357 | LD | 13030-L | 27513-LS-L | LOANER-SLX.LD.RETURN INTO INVENTORY(WAS SHIPPED FOR RMA01953) | LOANER RETURN. JS | RETESTED ALL THE DELIVERY DEVICES.JS | BCTechSupport@iridex.com | 2024-12-03 10:18 | Yes | No |  | No |  | No
SRV-RQ002099 | WI | 65600-08 | IQY72719C | cONVERT TO LNR, UIM Beeping activation when in standby on its own. Also changes screen randomly |  |  | BCTechSupport@iridex.com | 2024-12-03 10:55 | Yes | No |  | No |  | No
SRV-RQ002361 | PSD | 65600-01 | IQY71369C | CUST-IQ577, PSD, NEED EVALUATION, VERIFY SHIPPING ADDRESS BEFORE SHIPING | PM CHECK.JS | RETESTED ALL THE DELIVERY DEVICES NO PROBLEM FOUND.JS | BCTechSupport@iridex.com | 2024-12-03 11:45 | Yes | No |  | No |  | No
SRV-RQ002356 | PSD | 66245 | IQI91257C | CUST-G6.PSD.LOW POWER OUTPUT |  |  | BCTechSupport@iridex.com | 2024-12-03 13:03 | Yes | No |  | No |  | No
SRV-RQ002401 | PSD | 30903-H500 | 1380000170 | CUST-LIO.PSD.PM | Cannot test due to twisted abd broken assembly fiber optic, damaged extension | Replaced assembly fiber optic, extension cable, control stem, eccentric shaft, | BCTechSupport@iridex.com | 2024-12-03 13:43 | Yes | No |  | No |  | No
SRV-RQ002344 | PSD | 65900-H500 | K1752 | CUST-LIO.PSD.mirrors are out of alignment due to it falling | Cannot test due to Loose eye safety filters 577nm from oculars housing, |  | BCTechSupport@iridex.com | 2024-12-03 15:07 | Yes | No |  | No |  | No
SRV-RQ002362 | PSD | 87304 | 1101319494 | CUST-LIO, PSD, NEED EVALUATION, VERIFY SHIPPING ADDRESS BEFORE SHIPING | Cannot test due to twisted and broken assembly fiber optic, wrong foam LIO |  | BCTechSupport@iridex.com | 2024-12-04 07:50 | Yes | No |  | No |  | No
SRV-RQ002292 | PSD | 30903D-H500 | 1356295-D | CUST-LIO,PSD, OPTICS OUT OF ALIGNMENT, VERIFY SHIPPING BEFORE SHIPPING | Cannot test due to Missing assembly fiber optic, extension cable, cap fiber |  | BCTechSupport@iridex.com | 2024-12-04 11:29 | Yes | No |  | No |  | No
SRV-RQ002251 | PSD | 13030 | 29970-LS | CUST-SLx, PSD, NO ILLUMINATION | NO ILLUMINATION AS CUSTOMER COMPLAINT AND FOUND NO POWER | Reinstall aiming beam lens holder and Realign Laser head. RY | BCTechSupport@iridex.com | 2024-12-04 14:27 | Yes | No |  | No |  | No
SRV-RQ002377 | PSD | 65600-08 | IQY72171C | CUST-IQ577.PSD.PM | Customer return for PM check. SD | Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2024-12-05 07:34 | Yes | No |  | No |  | No
SRV-RQ002152 | PSI | 65777 | RC0471 | CUST-RC,PSI, PM (CAME WITH IQG81478C) | ENCODER WAS PARTLY WORKS AND FOUNDDIPLAY TO UIM WAS NOT | REPLACED ENCODER LCD DIPLAY AND GASKET RETESTED.JS | BCTechSupport@iridex.com | 2024-12-06 07:25 | Yes | No |  | No |  | No
SRV-RQ002141 | WD | 65900-H500 | K2155 | CUST-LIO,WD, BROKEN KNOBS | Cannot test due to Damaged extension cable, wrong and burnt halogen bulb, |  | BCTechSupport@iridex.com | 2024-12-06 07:37 | Yes | No |  | No |  | No
SRV-RQ002360 | PSD | 65600-08 | IQY72673C | Cust-577-PM-PSD&add shipping 1 wy (for box inbound) | Customer return for PM check. SD | Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2024-12-06 08:11 | Yes | No |  | No |  | No
SRV-RQ002318 | PSD | 65505 | REZZ15428 | CUST-SLA,PSD, LOW POWER AND NEED ALIGNMENT | Broken fiber, | Replaced fiber. Test clean and inspec. | BCTechSupport@iridex.com | 2024-12-06 13:07 | Yes | No |  | No |  | No
SRV-RQ002382 | PSD | 13152-H500 | 1362443 | CUST-LIO.PSD.MISALIGNMENT, NOT AIMING | Cannot test due to twisted and broken assembly fiber optic, bent control stem, |  | BCTechSupport@iridex.com | 2024-12-13 11:41 | Yes | No |  | No |  | No
SRV-RQ002422 | DCI | 70295 | YH14813TC | CUST- TXCELL. DCI. TAKES TOO LONG AND DOESN'T BOOT UP | UNABLE TO TESTED INTERFACE BOARD NOT WORKING PROPERLY. | REPLACED INTERFACE BOARD TESTED ALIGNED CALIBRARED CLEAN AND INSPECTED | BCTechSupport@iridex.com | 2024-12-13 09:10 | Yes | No |  | No |  | No
SRV-RQ002281 | SCD | 65900-H500 | 1362505 | cust-LIO.SCD.PM (Platinum contract) | Cannot test due to Twisted and burnt assembly fiber optic, broken control stem, |  | BCTechSupport@iridex.com | 2024-12-11 14:35 | Yes | No |  | No |  | No
SRV-RQ002281 | SCD | 65900-H500 | 1362505 | cust-LIO.SCD.PM (Platinum contract) | Cannot test due to Twisted and burnt assembly fiber optic, broken control stem, |  | BCTechSupport@iridex.com | 2024-12-11 14:31 | Yes | No |  | No |  | No
SRV-RQ002379 | PSD | 70294 | YEC13310TC | CUST-TxCell, PSD, led KEPT BLINKING |  |  | BCTechSupport@iridex.com | 2024-12-11 14:13 | Yes | No |  | No |  | No
SRV-RQ002427 | LD | 66245-L | IQI91310C-L | LOANER-IQ532.LD.RETURN INTO INVENTORY (CAME FROM CUST 14349) | Loaner return. SD | Retested all the delivery devives.SD | BCTechSupport@iridex.com | 2024-12-11 07:18 | Yes | No |  | No |  | No
SRV-RQ002426 | LD | 65500-L | IQG81577C-L | LOANER - IQ532.LD.RETURN INTO INVENTORY ( CAME FROM CUST 24983) |  |  | BCTechSupport@iridex.com | 2024-12-10 13:08 | Yes | No |  | No |  | No
SRV-RQ002404 | PSD | 31172-02 | GH11475LM | CUST-SLA, PSD, BAD FIBER not rcognized, just replaire month ago | LOOSE REGISTOR. | RE-TIGHTENING RESISTOR. TEST,CLEAN AND INSPECT. | BCTechSupport@iridex.com | 2024-12-10 13:07 | Yes | No |  | No |  | No
SRV-RQ002425 | LD | 65600-01-L | IQY71687C-L | LOANER - IQ577.RETURN INTO INVENTORY (CAME FROM CUST 23655) | Loaner return. SD | Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2024-12-10 08:03 | Yes | No |  | No |  | No
SRV-RQ002376 | PSD | 30903-H500 | K1466 | CUST.LIO.PSD.ILLUMINATON ISSUE | Cannot test due to Twisted and burnt assembly fiber optic, burnt halogen bulb, |  | BCTechSupport@iridex.com | 2024-12-09 14:51 | Yes | No |  | No |  | No
SRV-RQ002380 | PSD | 87304 | 1103738656 | CUST-LIO.PSD.MISALIGNMENT | Cannot test due to Twisted and broken assebly fiber optic and will need alignmen |  | BCTechSupport@iridex.com | 2024-12-09 12:37 | Yes | No |  | No |  | No
SRV-RQ002424 | LD | 65600-01-L | IQY71701C-L | LOANER -IQ577.RETURN INTO INVENTORY (WAS SHIPPED FOR CUST 23465 |  |  | BCTechSupport@iridex.com | 2024-12-09 12:26 | Yes | No |  | No |  | No
SRV-RQ001987 | PSD | 30905 | GL14698F30 | CUST-GL, PSD, INTERMITTENT CONNECT FIBER PROBLEM | Error 1,5,and 18. Console is not recognize the fiber CTRL board 30522 bad. |  | BCTechSupport@iridex.com | 2024-12-09 11:19 | Yes | No |  | No |  | No
SRV-RQ002348 | PSD | 65900-H500 | 1333595 | Cust-LIO, PSD, need PM and new headband, check fiberoptic, service and return | Cannot test due to Twistend and broken assembly fiber optic, wrong control stem, |  | BCTechSupport@iridex.com | 2024-12-09 11:07 | Yes | No |  | No |  | No
SRV-RQ002414 | PSD | 65600-01 | IQY71481C | CUST-IQ577.PSD.TXCELL WAS BLINKING AND CHECK IF THERE ANY ISSUE WITH IQ ALSO | TXCELL was blanking and check if there are any issue with IQ. SD | No proplems found on TXCELL and IQ durning tested process. SD | BCTechSupport@iridex.com | 2024-12-09 10:56 | Yes | No |  | No |  | No
SRV-RQ002103 | PSD | 30905 | GL14059F30 | CUST-GL.PSD.PM | PM CHECK, FOUND LOW POWER,NEED RECALIBRATION.JS | READJUSTED TEMPRATURE AND RECALIBRATED ALL THE DELIVERY DEVICES.JS | BCTechSupport@iridex.com | 2024-12-14 09:26 | Yes | No |  | No |  | No
SRV-RQ002237 | WD | 70300 | GH16519TC | CUST-TxCell,WD,LED BLINKING NON STOP | CANNOT DUPLICATED THE PROBLEM. | TESTED CALIBRATED CLEAN AND INSPECTED. | BCTechSupport@iridex.com | 2024-12-14 09:43 | Yes | No |  | No |  | No
SRV-RQ002335 | PSD | 13030 | 29960-LS | CUST SLX.PSD.PM |  |  | BCTechSupport@iridex.com | 2024-12-16 09:45 | Yes | No |  | No |  | No
SRV-RQ002322 | SCD | 66245 | IQI91631C | Cust-G6-GUI-SCD | upgrade software to gui.js | upgraded soft ware to gui and retested.js | BCTechSupport@iridex.com | 2024-12-16 10:22 | Yes | No |  | No |  | No
SRV-RQ002429 | LD | 13030-L | 29556-LS-L | LOANER-SLX.LD.RETURN INTO INVENTORY (CAME FROM CUST 15083) | Loaner return. SD | Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2024-12-16 10:43 | Yes | No |  | No |  | No
SRV-RQ002413 | PSI | 32002 | 18897H | CUST-HEAD-TX, PSI, LPA ERROR6 | Error 6 per customer's report. YO | Cust. Approved. DN | BCTechSupport@iridex.com | 2024-12-16 11:15 | Yes | No |  | No |  | No
SRV-RQ002412 | PSI | 32002 | 18451H | CUST. PSI. Laser Head | Laser power issue per customer reported. YO | Replaced 11053, 30264, 31720, 33039, 33047. YO | BCTechSupport@iridex.com | 2024-12-16 11:51 | Yes | No |  | No |  | No
SRV-RQ002421 | DCI | 70295 | YH15097TC | CUST-TXCELL.DCI.MOUNTING PLATE |  |  | BCTechSupport@iridex.com | 2024-12-16 12:55 | Yes | No |  | No |  | No
SRV-RQ002403 | PSI | 66245-1 | IQI92918C | CUST-G6, ( mfg.2019), PSI non-verified software not loading | UIM FAIL TO LOAD APPLICATION,BROKEN ENCODER NEED TO REPLACE | replaced encoder and retested all the delivery devices.js | BCTechSupport@iridex.com | 2024-12-16 12:56 | Yes | No |  | No |  | No
SRV-RQ002054 | PSD | 31215 | GL15911C30 | CUST-GLx, PSD, lOW POWER AT 200MW , NO BURN | Error 5 and Error 23. laser keep resting when firing power .Laser head Require | Realign laser head to manufacture GLX specification. Yo | BCTechSupport@iridex.com | 2024-12-16 13:06 | Yes | No |  | No |  | No
SRV-RQ002368 | PSD | 66245-1 | IQI92324C | Cust-G6.PSD. Per customer message, says the G6 is Glitching | CUSTOMER COMPLAINT GLICHING WHAT IS THE MEANING OF GLICHING | Realigned and calibration laser head. | BCTechSupport@iridex.com | 2024-12-16 14:09 | Yes | No |  | No |  | No
SRV-RQ002329 | PSD | 70300 | GH14541TC | CUST-TxCell, PSD, IN u-PUSE MODE GRID PATTERN SEEING ERRATIC SPEED SEE COMMENTS | NO PROBLEM WITH THE SPEED THAT'S THE NORMAL SPEED FOR THE NEW COMPUTER BOARD. | TESTED CLEAN AND INSPECTED. | BCTechSupport@iridex.com | 2024-12-17 13:04 | Yes | No |  | No |  | No
SRV-RQ002129 | PSI | 70292 | YEZ15217TC | CUST-TXCELL.PSI.PM PLEASE INVESTIGATE | PM CHECK INTERFACE BOARD AND USB CABLE NEED TO UPGRADED. | REPLACED INTERFACE BOARD USB CABLE TESTED ALIGNED CLEAN AND INSPECTED. | BCTechSupport@iridex.com | 2024-12-17 16:12 | Yes | No |  | No |  | No
SRV-RQ002295 | PSI | 65600-08 | IQY72129C | CUST-IQ577.STG.PM PLEASE INVESTIGATE | Customer return for PM check. SD | No problem found. Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2024-12-17 16:13 | Yes | No |  | No |  | No
SRV-RQ002370 | PSD | 70295 | YH13758TC | CUST-TXCELL, PSD , E05111 , | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING |  | BCTechSupport@iridex.com | 2024-12-18 06:34 | Yes | No |  | No |  | No
SRV-RQ002406 | PSI | 66245 | IQI91687C | CUST-G6.PSI.E05103 ERROR | Eo5103 dead batterry and power output low. SD | Realign and calibration laser head. RY | BCTechSupport@iridex.com | 2024-12-19 08:08 | Yes | No |  | No |  | No
SRV-RQ002396 | SCD | 66245-1 | IQI92344C | CUST-G6.SCD.PM/GUI | Broken encoder, SD | Replace encoder and retested all the delivery devices. SD | BCTechSupport@iridex.com | 2024-12-19 14:52 | Yes | No |  | No |  | No
SRV-RQ001506 | LD | 31214-L | GL14151F30-L | Loaner. LD. Return into inventory | Loaner returned for PM check up. PT | Tested all delivery device. Power read within spec. PT | BCTechSupport@iridex.com | 2024-12-23 09:20 | Yes | No |  | No |  | No
SRV-RQ002393 | PSD | 65500 | IQG81570C | CUST-IQ532.PSD.PM |  |  | BCTechSupport@iridex.com | 2024-12-23 12:25 | Yes | No |  | No |  | No
SRV-RQ002250 | PSD | 70292 |  | CUST- TXCELL.PSD.PM | BURNED FIBER AND CONTROL BOX NON-STOP BLINKING |  | BCTechSupport@iridex.com | 2024-12-23 13:59 | Yes | No |  | No |  | No
SRV-RQ002493 | PSD | 65600-01 | IQY71370C | CUST-IQ577.PSD.CAME WITH TXCELL EVALUATE IF ANY ISSUE |  |  | BCTechSupport@iridex.com | 2024-12-26 11:58 | Yes | No |  | No |  | No
SRV-RQ002463 | LD | 13030-L | 29730-LS-L | LOANER - SLX. LD . RETURN INTO INVENTORY (RETURN FROM CUST 15495) |  |  | BCTechSupport@iridex.com | 2024-12-26 13:01 | Yes | No |  | No |  | No
SRV-RQ002187 | PSD | 65600-01 | IQY71414C | CUST-IQ577.PSD.faulty TxCell and the IQ console (iq577) | LOW POWER AND ENCODER MAY NEED TO REPLACE HEAD SENT TO H.R.JS | -Replaced fiber detector board, realigned the laser head, recalibrated the unit. | BCTechSupport@iridex.com | 2024-12-27 06:28 | Yes | No |  | No |  | No
SRV-RQ002341 | CEM | 65600-01 | IQY71036C | IQ577.CEM.DAME PROJECT FOR JOAN | DEMO RETURN,FOUND BAD ENCODER,REPLACED ENCODER .js | Retestested all the delivery devices. JS | BCTechSupport@iridex.com | 2024-12-27 08:12 | Yes | No |  | No |  | No
SRV-RQ001532 | WD | 70292 | YEZ13080TC | CUST-TxCell,WD, LED keep blinking | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING, COMPUTER BOARD NOT WORKIN CAUSIN | REPLACED COMPUTER BOARD USB CABLE TESTED ALIGNED CALIBRATED CLEAN AND INSPECTED. | BCTechSupport@iridex.com | 2024-12-27 08:17 | Yes | No |  | No |  | No
SRV-RQ002473 | SCD | 66245-1 | IQI91999C | CUST-CYCLO-G6, SCD GOLD+, PM CHCKUP | PM CHECK. JS | RETESTED ALL THE DELIVERY DEVICES PER MFG SPEC.JS | BCTechSupport@iridex.com | 2024-12-27 09:57 | Yes | No |  | No |  | No
SRV-RQ002440 | PSD | 70295 | YH13035TC | CUST-TxCell, PSD, LED KEEP BLINKING | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING COMPUTER BOARD |  | BCTechSupport@iridex.com | 2024-12-28 09:17 | Yes | No |  | No |  | No
SRV-RQ002417 | CEM | 12731 | 1391 | GL, SL, G6 CEM,Calibration and PM | Calibration Due. | Retested. RY | BCTechSupport@iridex.com | 2024-12-30 06:28 | Yes | No |  | No |  | No
SRV-RQ002452 | PSD | 65500 | IQG81571C |  | PM Check. | Retested all the delivery devices as MFG spec.js | BCTechSupport@iridex.com | 2024-12-31 07:12 | Yes | No |  | No |  | No
SRV-RQ002338 | PSD | 65600-01 | IQY71526C | CUST-IQ577.PSD.PM | Power output low. SD | -Replaced fiber focus lens, realigned the laser head, recalibrated the unit. | BCTechSupport@iridex.com | 2024-12-31 07:14 | Yes | No |  | No |  | No
SRV-RQ002417 | CEM | 66245 | IQI91364C | GL, SL, G6 CEM,Calibration and PM | Calibration and PM | Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2024-12-31 08:39 | Yes | No |  | No |  | No
SRV-RQ002303 | PSD | 31172-01 | GH11025LM | cust-SLA.PSD.UNABLE TO CENTER |  |  | BCTechSupport@iridex.com | 2024-12-31 08:56 | Yes | No |  | No |  | No
SRV-RQ002319 | PSD | 70292 | YEZ13630TC | CUST-TXCELL, PSD, PM CHECK / NOW HAS E0511 ERROR CODE | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING COMPUTER BOARD NOT |  | BCTechSupport@iridex.com | 2024-12-31 09:58 | Yes | No |  | No |  | No
SRV-RQ002388 | PSD | 65600D-01 | IQY71455C-D | CUST-IQ577.PSD.JUMPY ENCODER | Encoder is borken and power is low .Need replace encoder board #65297 and laser |  | BCTechSupport@iridex.com | 2024-12-31 12:31 | Yes | No |  | No |  | No
SRV-RQ002504 | PSD | 70292 | YEZ13429TC | CUST-TxCell.PSD.PM | CONTROL BOX NON-STOP BLINKING COMPUTER BOARD NOT WORKING |  | BCTechSupport@iridex.com | 2024-12-31 12:45 | Yes | No |  | No |  | No
SRV-RQ002419 | PSD | 30903-H500 | 1341771 | CUST-LIO, PSD, CHECK FOR ALIGNMENT , NO CHARGE fIBEROPTIC ,JUST REPLACED | Cannot test dur to broken assembly fiber optic, bent control stem, burnt halogen |  | BCTechSupport@iridex.com | 2024-12-31 15:09 | Yes | No |  | No |  | No
SRV-RQ002416 | PSD | 70292D | YEZ14143TC-D | CUST-TxCell, PSD, LEDV KEEP BLINKING | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING |  | BCTechSupport@iridex.com | 2025-01-02 08:52 | Yes | No |  | No |  | No
SRV-RQ002509 | CEM | 65600D-01 | IQY71036C-D | IQ577.CEM.DAME PROJECT FOR JOAN | DAME project request per Joan. JS | Replaced encoder and retested all delivery devices. JS | BCTechSupport@iridex.com | 2025-01-02 09:30 | Yes | No |  | No |  | No
SRV-RQ002392 | PSD | 66245 | IQI91388C | CUST-G6.PSD.PM | Customer return for PM check. SD | Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2025-01-02 10:02 | Yes | No |  | No |  | No
SRV-RQ002350 | PSD | 30903-H500 | K1723 | CUST-LIO, PSD, bROKEN NO OTHER PROBLEM REPORTED, VERIFY SHIPPING ADRESS | Cannot test due to Twsited and burnt assembly fiber optic, damaged extension |  | BCTechSupport@iridex.com | 2025-01-02 10:33 | Yes | No |  | No |  | No
SRV-RQ002437 | PSD | 30903-H500 | L3417 | Cust-LIO PSD, mirror needs replacment per customer, proivde an estimate before work | T-slot adapter and will need alignment. |  | BCTechSupport@iridex.com | 2025-01-02 11:24 | Yes | No |  | No |  | No
SRV-RQ002492 | SCD | 70295 | YH13311TC | CUST-TxCell, SCD-P, LED KEEP BLINKING | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING |  | BCTechSupport@iridex.com | 2025-01-02 12:19 | Yes | No |  | No |  | No
SRV-RQ002450 | PSD | 65900-H500 | K1017 | CUST-LIO, PSD, WHITE LIGHT CIRCLE CUT OFF, NO AIMBEAM | Cannot test dur to burnt assembly fiber optic, burnt haloge bulb, viewing |  | BCTechSupport@iridex.com | 2025-01-02 15:11 | Yes | No |  | No |  | No
SRV-RQ002454 | PSD | 65515-H500 | 1362968 | CUST-LIO, PSD, LOW POWER AND EVALUATION NEEDED | Cannot test du to Twisted and burnt assembly fiber optic, bent control stem, |  | BCTechSupport@iridex.com | 2025-01-02 14:28 | Yes | No |  | No |  | No
SRV-RQ002458 | PSD | 30903-H500 | K1442 | Cust-LIO, PSD, Check LIO for a broken mirror | Cannot test due to Twisted and burnt assembly fiber optic, damaged exension |  | BCTechSupport@iridex.com | 2025-01-03 09:33 | Yes | No |  | No |  | No
SRV-RQ002378 | PSD | 66245-1 | IQI92315C | CUST-G6.PSD.PM | Broken encoder | Replace encoder and retested all the delivery devices. | BCTechSupport@iridex.com | 2025-01-03 08:39 | Yes | No |  | No |  | No
SRV-RQ002464 | LD | 65600-01-L | IQY71716C-L | LOANER-IQ577. LD . RETURN INTO INVENTORY ( RETURN FROM CUST 24805) | PHOTOCELL MISMATCH. HR.JS | -Replaced the PCBA adapter , recalibrated the unit. | BCTechSupport@iridex.com | 2025-01-03 05:30 | Yes | No |  | No |  | No
SRV-RQ002319 | PSD | 65600-01 | IQY71459C | CUST-IQ577, PSD, PM CHECK / NOW HAS E0511 ERROR CODE | PM CHECK.JS |  | BCTechSupport@iridex.com | 2025-01-03 05:22 | Yes | No |  | No |  | No
SRV-RQ002503 | LD | 66245-2-L | IQI92329C-L | LOANER-LD.RETURN INTO INVENTORY (CAME FROM CUST 12194) | Loaner return.SD | Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2025-01-03 10:35 | Yes | No |  | No |  | No
SRV-RQ002469 | SCD | 65600-01 | IQY71458C | CUST-IQ577, SCD , LOOSE KNOBS, NEED EVALUATION | Has jumpy encoder and port 2 intermittengly recognise after replaced both |  | BCTechSupport@iridex.com | 2025-01-03 11:24 | Yes | No |  | No |  | No
SRV-RQ002458 | PSD | 30903-H500 | K1442 | Cust-LIO, PSD, Check LIO for a broken mirror | No Problem found just need some adjusment. | Refocus laser spot size beam centered. | BCTechSupport@iridex.com | 2025-01-03 11:51 | Yes | No |  | No |  | No
SRV-RQ002468 | PSD | 30903-H500 | K1426 | CUST-LIO, PSD, LOW POWER , NEED ALIGNMENT | Cannot test due to Twisted and burnt assembly fiber optic, damaged extension |  | BCTechSupport@iridex.com | 2025-01-03 12:03 | Yes | No |  | No |  | No
SRV-RQ002390 | PSD | 65505 | REZZ12445 | CUST-SLA.PSD.MISALIGNMENT EXPOSE WIRES |  |  | BCTechSupport@iridex.com | 2025-01-03 12:23 | Yes | No |  | No |  | No
SRV-RQ002439 | WD | 30903-H500 | 1344082 | CUST-LIO, WD, INTERMITTENT CONNECTIVITY PROBLEM JUST REPAIRED | No prblem found. | Beam centere and refocus laser spot size. | BCTechSupport@iridex.com | 2025-01-03 15:39 | Yes | No |  | No |  | No
SRV-RQ002397 | SCD | 65500 | IQG81586C | CUST-IQ532, SCD, NEED EVALUATION PM | PM CHECK FOOT SWITCH CAME WITH UNIT.JS | RETESTED ALL THE DELIVERY DFEVICES AND FOOT SWITCH WORKS FINE.JS | BCTechSupport@iridex.com | 2025-01-03 13:37 | Yes | No |  | No |  | No
SRV-RQ002453 | PSD | 31172-01 | GH16247LM | CUST-SLA, PSD, AIMBEAM NOT FOCUSING, NEED ALIGNMENT |  |  | BCTechSupport@iridex.com | 2025-01-03 13:20 | Yes | No |  | No |  | No
SRV-RQ002485 | SD | 66245-1 | IQI92168C | CUST-CYCLO-G6, PSD , LOW POWER PLEASE EVLUATE | low power as customer complaint sent to head room.JS | Realign and calibration laser head. RY | BCTechSupport@iridex.com | 2025-01-06 09:06 | Yes | No |  | No |  | No
SRV-RQ002461 | PSD | 66245-1 | IQI92448C | Cust-G6.PSD.GUI Upgrade |  |  | BCTechSupport@iridex.com | 2025-01-06 08:32 | Yes | No |  | No |  | No
SRV-RQ002324 | WD | 66245-1 | IQI93786C | Convert to Inventory-G6-WD-DON'T SHIP | Convert to inventory G6 | Convert to Demo and retested all the delivery devices. SD | BCTechSupport@iridex.com | 2025-01-06 07:38 | Yes | No |  | No |  | No
SRV-RQ002418 | PSD | 65600-01 | IQY71789C | CUST-IQ577.PSD.UIM ERRROR APPLICATION | Power output low .SD |  | BCTechSupport@iridex.com | 2025-01-06 14:27 | Yes | No |  | No |  | No
SRV-RQ002480 | SD | 65500 | IQG81664C | CUST-IQ532.PSD.PM |  |  | BCTechSupport@iridex.com | 2025-01-07 08:35 | Yes | No |  | No |  | No
SRV-RQ002400 | PSD | 31214 | GL11125 | CUST GL.PSD.PM | Temperature of laser head is not stable ,sound run up and down very loud . also | No part and no station to repair. YO | BCTechSupport@iridex.com | 2025-01-07 06:48 | Yes | No |  | No |  | No
SRV-RQ002476 | SCD | 70292 | YEZ13039TC | CUST-TxCell, SCD GOLD+, LED KEEP BLINKING | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING COMPUTER BOARD NOT WORKING |  | BCTechSupport@iridex.com | 2025-01-07 12:28 | Yes | No |  | No |  | No
SRV-RQ002487 | SD | 13030 | 30445-LS | CUST-CYCLO-SLx, WD , PM CHECKUP |  |  | BCTechSupport@iridex.com | 2025-01-08 06:48 | Yes | No |  | No |  | No
SRV-RQ001407 | PSD | 30905 | GL14326F30 | CUST-GL PSD, OcuLight GL has a smell and the indirect illumination hard to insert and out | Console was burnt and smell . found LIO board burnt . LIO Jack too tight . | Replaced heaterboard assemply. YO | BCTechSupport@iridex.com | 2025-01-09 13:10 | Yes | No |  | No |  | No
SRV-RQ002437 | PSD | 30903-H500 | L3417 | Cust-LIO PSD, mirror needs replacment per customer, proivde an estimate before work | T-slot adapter, mirror turn damaged coating, extension cable read low power |  | BCTechSupport@iridex.com | 2025-01-09 08:23 | Yes | No |  | No |  | No
SRV-RQ002465 | SI | 65600-01 | IQY71389C | CUST-IQ577, PSI , Error E6001, LOW POWER ACTUAL | low power as customer complaint send to head room js |  | BCTechSupport@iridex.com | 2025-01-08 14:09 | Yes | No |  | No |  | No
SRV-RQ002087 | SI | 65500-08 | IQG82080C | CUST-IQ532, WI, per caller there is no aming beam?! Please verify then repair | NO AIMING BEAM AS CUSTOMER COMPLAINT AND VERY POW POWER | -Replaced Aiming beam diode, Assy cable photocell, realigned the laser head, | BCTechSupport@iridex.com | 2025-01-14 12:55 | Yes | No |  | No |  | No
SRV-RQ002524 | SD | 13030 | 29572-LS | CUST-SLx, PSD, LPA , PLEASE EVALUATE | Need calibration | Recalibrated all the delivery devices as per MFG spec.js | BCTechSupport@iridex.com | 2025-01-13 12:22 | Yes | No |  | No |  | No
SRV-RQ001464 | LD | 30658-L | GL14031C30-L | LD. Loaner GLx. Return into inventory | Power shuts off when turn on for few minute .found power supply board is bad . | This console old .Power supply board bad .No parts for repair.Require SCRAP .CT | BCTechSupport@iridex.com | 2025-01-13 10:14 | Yes | No |  | No |  | No
SRV-RQ002495 | PSD | 31172-01 | GH10243LM | CUST-SLA. PSD, U-MANUPULATOT ASSAY BROKEN |  |  | BCTechSupport@iridex.com | 2025-01-10 13:20 | Yes | No |  | No |  | No
SRV-RQ001682 | PSD | 30465 | GL12991B | CUST-GL, PSD, LOOSE SMARTKEY CONNECTOR | Loose smartkey connector , Error 5 and Error 23. Laser head require Eval ct | Retigh smarkkey connector . Recalibrated all delivery devices .Retest and burn | BCTechSupport@iridex.com | 2025-01-10 08:51 | Yes | No |  | No |  | No
SRV-RQ002350 | PSD | 30903-H500 | K1723 | CUST-LIO, PSD, bROKEN NO OTHER PROBLEM REPORTED, VERIFY SHIPPING ADRESS | Cannot test due to Twsited and burnt assembly fiber optic, damaged extension |  | BCTechSupport@iridex.com | 2025-01-15 10:50 | Yes | No |  | No |  | No
SRV-RQ002350 | PSD | 30903-H500 | K1723 | CUST-LIO, PSD, bROKEN NO OTHER PROBLEM REPORTED, VERIFY SHIPPING ADRESS | Cannot test due to Twsited and burnt assembly fiber optic, damaged extension |  | BCTechSupport@iridex.com | 2025-01-15 11:05 | Yes | No |  | No |  | No
SRV-RQ002562 | SD | 12733 | 1189-VET | CUST-VET, PSD, EROOR 15 |  |  | BCTechSupport@iridex.com | 2025-01-15 11:17 | Yes | No |  | No |  | No
SRV-RQ002460 | PSD | 70295 | YH13230TC | Cust-TxCell.PSD. E05111 (Green light blinkng) | CONTROL BOX NON-STOP BLINKING COMPUTER BOARD CAUSING THE PROBLEM. |  | BCTechSupport@iridex.com | 2025-01-15 17:29 | Yes | No |  | No |  | No
SRV-RQ002554 | WD | 65900-H500 | K2342 | CUST-LIO, WD, FRONT LENSE LOOSE INSIDE | Cannot test due to delivery mirror loose, damaged exension cable, beam clipping | Replaced assembly fiber optic, reglue delivery mirror, restored extension cable, | BCTechSupport@iridex.com | 2025-01-16 13:43 | Yes | No |  | No |  | No
SRV-RQ002468 | PSD | 30903-H500 | K1426 | CUST-LIO, PSD, LOW POWER , NEED ALIGNMENT | Cannot test due to Twisted and burnt assembly fiber optic, damaged extension |  | BCTechSupport@iridex.com | 2025-01-17 10:44 | Yes | No |  | No |  | No
SRV-RQ002534 | PSI | 65600-01 | IQY71149C | CUST-IQ577, PSI, ERROR CODE TWICE- NO OTHER PROBLEM STATED | Console was drop down .Found front panel #65892,t Top chasiss handle broken . |  | BCTechSupport@iridex.com | 2025-01-17 11:14 | Yes | No |  | No |  | No
SRV-RQ002515 | PSD | 65777 | RC0336 | CUST-REMOTE, PSD, SCREEN GOES BLANK, NOT WORKING | UIM failed to load. | Reconnect SD card and retested all the devices. | BCTechSupport@iridex.com | 2025-01-17 14:07 | Yes | No |  | No |  | No
SRV-RQ002417 | CEM | 30465D | GL12714B-D | GL, SL, G6 CEM,Calibration and PM | Low Aiming Beam and ferruler stripped . Laser head require Eval. CT | Realign laser head to manufacture GL2 thin plate specification. Yo | BCTechSupport@iridex.com | 2025-02-03 07:30 | Yes | No |  | No |  | No
SRV-RQ002594 | PSD | 12733 | 1195-VET | CUST-DIOVET, PSD, EMERGENCY STOP MESSAGE | Power output low. SD | RECAL/RETESTED all the delivery devices. SD | BCTechSupport@iridex.com | 2025-01-31 11:02 | Yes | No |  | No |  | No
SRV-RQ002544 | SCD | 66245-1 | IQI92985C | CUST-CYCLO-G6, SCD, PM CHECKUP | LOW POWER ON PORT 1 . Laser head require Eval, CT | Realign and calibration laser head. RY | BCTechSupport@iridex.com | 2025-01-30 10:57 | Yes | No |  | No |  | No
SRV-RQ002505 | PSD | 66245 | IQI91550C | CUST-G6.PSD.LOW POWER | Low power output.SD | Realign and calibration laser head. RY. | BCTechSupport@iridex.com | 2025-01-30 10:19 | Yes | No |  | No |  | No
SRV-RQ002505 | PSD |  |  | CUST-G6.PSD.LOW POWER | Low power output.SD | Realign and calibration laser head. RY.Replace PCBA fiber detect board. SD | BCTechSupport@iridex.com | 2025-01-30 10:19 | Yes | No |  | No |  | No
SRV-RQ002441 | SD | 65600-08 | IQY72847C | CUST-IQ577 OBF E06004, WD, (DELIVER TO PRITAM only). Need confirmation of the error before repairs |  |  | BCTechSupport@iridex.com | 2025-01-30 10:16 | Yes | No |  | No |  | No
SRV-RQ002546 | SD | 66245-1 | IQI92479C | CUST-CYCLO-G6, PSD, PM CHECKUP |  |  | BCTechSupport@iridex.com | 2025-01-30 08:04 | Yes | No |  | No |  | No
SRV-RQ002563 | DD | 65515-H500 | K2448 | SERVICE, DD | No Problems Found, PM LIO, Replace Bulb and Dust Cap w/Lanyard, | PM LIO, Replaced Bulb and Dust Cap w/ Lanyard, Tighted all Screws, | BCTechSupport@iridex.com | 2025-01-29 08:08 | Yes | No |  | No |  | No
SRV-RQ002613 | SD | IQ532-SYSTEM-D | IQG81642C-D | demo.DD.RETURN TO INVENTORY |  |  | BCTechSupport@iridex.com | 2025-01-29 06:49 | Yes | No |  | No |  | No
SRV-RQ002598 | WI | 65600-01-L | IQY71302-L | IQ + TxCell (2014) DAME project (Joan S.) Please check ( URGENT ) and inform her when ready | Demo returned. SD | Retested all delivery devices. SD | BCTechSupport@iridex.com | 2025-01-28 08:50 | Yes | No |  | No |  | No
SRV-RQ002598 | WI | 70295D | YH13056D | IQ + TxCell (2014) DAME project (Joan S.) Please check ( URGENT ) and inform her when ready | Demo returned. SD | Retested all delivery devices. SD | BCTechSupport@iridex.com | 2025-01-28 08:50 | Yes | No |  | No |  | No
SRV-RQ002417 | CEM | 33005 | GLX18727C | GL, SL, G6 CEM,Calibration and PM | PM require. CT | Replaced controller pcba. CT | BCTechSupport@iridex.com | 2025-01-28 07:49 | Yes | No |  | No |  | No
SRV-RQ002614 | SD | 31808D | 112390F-D | DEMO RETURN. DD. BACK TO INVENTORY |  |  | BCTechSupport@iridex.com | 2025-01-28 07:27 | Yes | No |  | No |  | No
SRV-RQ002530 | PSD | 70292 | YEZ13141TC | CUST-TxCell, PSD, LED KEEP BLINKING | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING |  | BCTechSupport@iridex.com | 2025-01-27 10:48 | Yes | No |  | No |  | No
SRV-RQ002586 | PSI | 70292D | YEZ14706TC-D | CUST-TxCell, PSI,LED BLINKING INTERMITTENTLY, s | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING |  | BCTechSupport@iridex.com | 2025-01-27 10:42 | Yes | No |  | No |  | No
SRV-RQ002590 | SD | 12733 | 767-VET | CUST-DIOVET, PSD, EMERGENCY STOP ERROR |  |  | BCTechSupport@iridex.com | 2025-01-27 08:42 | Yes | No |  | No |  | No
SRV-RQ002598 | WI | 65600-01-L | IQY71302-L | IQ + TxCell (2014) DAME project (Joan S.) Please check ( URGENT ) and inform her when ready | Loaner return. SD | Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2025-01-24 13:13 | Yes | No |  | No |  | No
SRV-RQ002598 | WI | 70295D | YH13056D | IQ + TxCell (2014) DAME project (Joan S.) Please check ( URGENT ) and inform her when ready | Loaner return. SD | Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2025-01-24 13:13 | Yes | No |  | No |  | No
SRV-RQ002535 | PSD | 30903-H500 | 1362369 | CUST-LIO,PSD, MIRROR FELL OFF INSDE | Cannot test dur to delivery mirror loosen from flexure support and scratched, |  | BCTechSupport@iridex.com | 2025-01-23 12:18 | Yes | No |  | No |  | No
SRV-RQ002535 | PSD | 30903-H500 | 1362369 | CUST-LIO,PSD, MIRROR FELL OFF INSDE | Cannot test dur to delivery mirror loosen from flexure support and scratched, |  | BCTechSupport@iridex.com | 2025-01-22 11:16 | Yes | No |  | No |  | No
SRV-RQ002387 | SD | 65600-01 | IQY71600C | CUST-IQ577.PSD.E05111, VERFY SHIPTO |  |  | BCTechSupport@iridex.com | 2025-01-22 07:42 | Yes | No |  | No |  | No
SRV-RQ002457 | PSD | 31808 | 112045F | CUST-FOOTSWITCH, PSD , DEAD BATTERIES |  |  | BCTechSupport@iridex.com | 2025-01-20 14:20 | Yes | No |  | No |  | No
SRV-RQ001992 | PSD | 31808 | 112594F | CUST-FOOTSWITCH, PSD, NO COMMUNICATION, RED LIGHT FLASHES | No communication . CT | Need send back thid footswitch to supplier for reapir CT | BCTechSupport@iridex.com | 2025-01-20 13:40 | Yes | No |  | No |  | No
SRV-RQ002514 | SD | 65500-08 | IQG81818C | CUST-IQ532, PSD, EVALUATE FULLY NO PROBLEM STATED, NOT WORKING | Damaged display unit not boot up, need to replace display at least have no |  | BCTechSupport@iridex.com | 2025-01-20 10:01 | Yes | No |  | No |  | No
SRV-RQ001396 | DD | 65777D | RC10274 | DEMO REMOTE. DD. RETURN TO INVENTORY | Demo return. SD | Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2025-01-20 08:36 | Yes | No |  | No |  | No
SRV-RQ002693 | WI | 70295 | YH15441TC | CUST-TXCELL.WI.E09505 Again after serviced in July'23 | CONTROL BOX NON-STOP BLINKING INTERFACE BOARD CAUSING THE PROBLEM. | REPLACED INTERFACE BOARD TESTED CALIBRATED CLEAN AND INSPECTED. | BCTechSupport@iridex.com | 2025-02-28 11:25 | Yes | No |  | No |  | No
SRV-RQ002519 | PSD | 30266 | GH5120 | CUST-SLA, PSD, Broken PARTS, CUSTOMER KNOW 30 YEAR OLD MAY BE SOME PARTS NOT AVAIL |  |  | BCTechSupport@iridex.com | 2025-02-28 10:41 | Yes | No |  | No |  | No
SRV-RQ002536 | PSD | 65515-H500 | K1925 | CUST-LIO, PSD, MISSING LENS | Cannot test due to missig eye safety filter 532 nm, burnt halogen bulb, missing |  | BCTechSupport@iridex.com | 2025-02-28 09:17 | Yes | No |  | No |  | No
SRV-RQ002733 | SD | 65600-08 | IQY72519C | CUST-IQ577. PSD. Power Surg happened no display | NOT BOOT UP,BECAUSE SCM PCBA BLOW UP NEED TO REPLACE.JS | REPLACED SCM PCBA AND RETESTED ALL THE DELIVERY DEVICES. JS | BCTechSupport@iridex.com | 2025-02-28 08:10 | Yes | No |  | No |  | No
SRV-RQ002561 | PSD | 32000-1 | TX1011REVB | CUST-TX, PSD, ERROR 38 POWER SUPPLY | ERROR 2 AND ERORR 38 . P/S BOARD IS BAD . RETEST FOUNG ENT HIGH DAC OUT OF RANGE | Realign laser head to manufacture TX specification.Yo | BCTechSupport@iridex.com | 2025-02-28 06:06 | Yes | No |  | No |  | No
SRV-RQ002622 | PSI | 66245-1 | IQI93185C | CUST-G6, E06006, PSI ( oler G6 and its PC is out of range) | Power output out of range | Realign and calibration Laser head. RY | BCTechSupport@iridex.com | 2025-02-27 06:51 | Yes | No |  | No |  | No
SRV-RQ002526 | WI | 33004-1 | GL41584C | Cust-GL, WI, LPA < 300mw and max power output is 70mw. | Error 5 . power keep reset with message "LOW POWER ACTUAL" | Realign laser head to manufacture GL NEO specification. Yo | BCTechSupport@iridex.com | 2025-02-27 06:33 | Yes | No |  | No |  | No
SRV-RQ002632 | SD | 65600-01 | IQY71352C | CUST-IQ577, SCD, BROKEN DROP TREAT to STD. BY | CUSTOMER DROP THE UNIT AND MISALIGNED FIBER ADAPTER AFTER ALIGNED FOUND |  | BCTechSupport@iridex.com | 2025-02-26 14:46 | Yes | No |  | No |  | No
SRV-RQ002720 | PSD | 65900-H500 | K2017 | CUST-LIO , PSD, BROKEN DROP TREAT to STD. BY | twisted assembly fiber optic, broken over band Omega, will need adjustment and |  | BCTechSupport@iridex.com | 2025-02-26 12:49 | Yes | No |  | No |  | No
SRV-RQ002698 | PSI | 70292 | YEZ15988TC | CUST-TxCELL. PSI. Not communicating with IQ577 IQY72670C | CANNOT DUPLICARED THE PROBLEM BUT OUT PUTLENS IS LOOSE NEED TO BE GLUE. | RE GLUE OUT PUTLENS TESTED ALIGNED CALIBRATED CLEAN AND INSPECTED. | BCTechSupport@iridex.com | 2025-02-26 12:31 | Yes | No |  | No |  | No
SRV-RQ002640 | SCD | 65600-08 | IQY72857C | CUST-IQ577, SCD, DROPPING T TO STDBY | Power output low. SD | -Realigned the laser head, recalibrated the unit. | BCTechSupport@iridex.com | 2025-02-25 14:04 | Yes | No |  | No |  | No
SRV-RQ002527 |  | 65600-08 | IQY72700C | CUST-IQ577, WD, SHUT DOWN ITERMITTENTLY | INTERMITTENGLY SHUTS OFF CUSTOMER COMPLAINT, HAVE NOT SEEN DURING PROCSS.JS | RETESTED ALL THE DELIVERY DEVICES AND PRCOESS BURNED COMPLETED.JS | BCTechSupport@iridex.com | 2025-02-24 10:50 | Yes | No |  | No |  | No
SRV-RQ002610 | PSD | 66120 | IQHG81008C | CUST-IQHG81008C, PSD, PM needed | System temperature out of range.sd | -Realigned the laser head, recalibrated the unit. | BCTechSupport@iridex.com | 2025-02-21 13:47 | Yes | No |  | No |  | No
SRV-RQ002545 | SD | 32000-4 | TX18907C | CUST-TX, PSD, LOW POWER MESSAGE, call for adress before shipping | DUPLICATED CUSTOMER COMPLAINT AS MENTION HEAD SENT TO HEAD |  | BCTechSupport@iridex.com | 2025-02-21 13:44 | Yes | No |  | No |  | No
SRV-RQ002522 | PSD | 32000-1 | TX18149C | CUST-TX, Error 38, PSD, | Error 2 and Error 38 .Power keep reset when firing power and clicking noise . |  | BCTechSupport@iridex.com | 2025-02-21 11:06 | Yes | No |  | No |  | No
SRV-RQ002653 | PSD | 65900-H500 | K1384 | CUST-LIO, PSD, BROKEN KNOBS AND UNABLE TO CENTER IMAGE | Cannot test due to control knobs do not move up and down, loose housing |  | BCTechSupport@iridex.com | 2025-02-21 10:51 | Yes | No |  | No |  | No
SRV-RQ002517 | SD | 65500 | IQG81442C | CUST-IQ532. PSD. Evaluate jumpy encoder | JUMPY ENCODER AS CUSTOMER COMPLAINT NEED TO REPLACE BEFORE INVESTIGATE.JS |  | BCTechSupport@iridex.com | 2025-02-21 09:27 | Yes | No |  | No |  | No
SRV-RQ002669 | PSD | 30903-H500 | K2104 | CUST-LIO, PSD, DOUBLE VISION | Cannot test due to Loose housing illumination and missing phillip screw, loose |  | BCTechSupport@iridex.com | 2025-02-21 09:05 | Yes | No |  | No |  | No
SRV-RQ002669 | PSD | 30903-H500 | K2104 | CUST-LIO, PSD, DOUBLE VISION | Cannot test due to Loose housing illumination and missing phillip screw, loose |  | BCTechSupport@iridex.com | 2025-02-21 09:01 | Yes | No |  | No |  | No
SRV-RQ002669 | PSD | 30903-H500 | K2104 | CUST-LIO, PSD, DOUBLE VISION | Cannot test due to Loose housing illumination and missing phillip screw, loose |  | BCTechSupport@iridex.com | 2025-02-21 09:00 | Yes | No |  | No |  | No
SRV-RQ002506 | PSD | 31215 | GL15638C30 | CUST-GLX.PSD.LOW POWER *HOT* | Error 5 and error 24 . Head board was burn can't firing power . Need replace | Realign laser head to manufacture GL 3rd cooler Specification. Yo | BCTechSupport@iridex.com | 2025-02-20 11:19 | Yes | No |  | No |  | No
SRV-RQ002618 | SI | 66245-1 | IQI93184C | CUST-G6, PSI, ' E05103 | UNIT CAME BECAUSE ERROR E05103 HAVE NOT DUPLICATED. FOOT SWITCH CAME WITH | Since the error relates to the battery, replaced the battery. JS | BCTechSupport@iridex.com | 2025-02-20 10:58 | Yes | No |  | No |  | No
SRV-RQ002709 | WI | 65600-01-L | IQY71302C-L | IQ577 Laser. WI. Back into inventory | Demo returned. SD | Retested all delivery devices. SD | BCTechSupport@iridex.com | 2025-02-20 10:42 | Yes | No |  | No |  | No
SRV-RQ002694 | PSI | 65600-08 | IQY72670C | CUST-IQ577. PSI. Not communicating with TxCell YEZ15988TC | Not communicating with TxCell YEZ15988TC.SD | No problem found for this console . Retested al the delivery devices. sd | BCTechSupport@iridex.com | 2025-02-20 08:11 | Yes | No |  | No |  | No
SRV-RQ002537 | PSD | 70292 | YEZ13712TC | cust-TxCell, psd, Eo5111 fiberoptic cable compromised | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING |  | BCTechSupport@iridex.com | 2025-02-20 07:53 | Yes | No |  | No |  | No
SRV-RQ002637 | PSD | 70292 | YEZ13179TC | CUST-TXCELL, PSD, LED KEEP BLINKING, AUX NOT MESSAGE | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING |  | BCTechSupport@iridex.com | 2025-02-20 07:46 | Yes | No |  | No |  | No
SRV-RQ002644 | PSD | 70292 | YEZ13264TC | CUST-TXCEL, PSD, led BLINKING, vERIFY SHIP TO ADRESS BEFORE SHIPPING | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING |  | BCTechSupport@iridex.com | 2025-02-19 11:19 | Yes | No |  | No |  | No
SRV-RQ002479 | PSI | 70292 | YEZ14762TC | CUST-TxCell,PSI, INTERMITTENTLY EO500,EO5111, REPLACED SCANNER BOARD AND BATTRY AND BIOS DONE | CANNOT DUPLICATE THE PROBLEM . | TESTED CALIBRATED ALIGNED CLEAN AND INSPEDTED | BCTechSupport@iridex.com | 2025-02-19 09:27 | Yes | No |  | No |  | No
SRV-RQ002549 | PSD | 30903-H | 1170314 | CUST-LIO, PSD, lOW ILLUMINATIOM LIGHT | Cannot test due to bunrt assembly fiber optic, bent control stem, burnt halogen, |  | BCTechSupport@iridex.com | 2025-02-18 15:38 | Yes | No |  | No |  | No
SRV-RQ002549 | PSD | 30903-H | 1170314 | CUST-LIO, PSD, lOW ILLUMINATIOM LIGHT | Cannot test due to bunrt assembly fiber optic, bent control stem, burnt halogen, |  | BCTechSupport@iridex.com | 2025-02-18 14:51 | Yes | No |  | No |  | No
SRV-RQ002679 | PSD | 65600-01 | IQY71563C | CUST-IQ577, PSD, PORT ONE NOT WORKING | Port 1 is not working . Can't firing . found Fibert detect board is bad #65451 | Replaced fiber Detect board #65451 . Recalibrated all delivery devices ,retest | BCTechSupport@iridex.com | 2025-02-14 12:53 | Yes | No |  | No |  | No
SRV-RQ002520 | PSD | 70297 | EZZ15431TC | CUST-TXCELL, PSD, NO AIM BEAM , LOW POWER, shiping chg 93.38 | UNABLE TO TESTED FIBER IS BROKEN. |  | BCTechSupport@iridex.com | 2025-02-14 10:02 | Yes | No |  | No |  | No
SRV-RQ002264 | PSD | 30903-H500 | 1361749 | CUST-LIO,PSD, FRONT MIRROR BROKEN | Cannot test due to Twisted and burnt assembly fiber optic, burnt halogen bulb, |  | BCTechSupport@iridex.com | 2025-02-14 09:19 | Yes | No |  | No |  | No
SRV-RQ002679 | PSD | 65600-01 | IQY71563C | CUST-IQ577, PSD, PORT ONE NOT WORKING | Port 1 is not working . Can't firing . found Fibert detect board is bad #65451 | Replaced fiber Detect board #65451 . Recalibrated all delivery devices ,retest | BCTechSupport@iridex.com | 2025-02-13 13:59 | Yes | No |  | No |  | No
SRV-RQ002385 | PSD | 32000-1 | TX18309C | CUST-Tx, PSD, error 2, please evaluate and provide an estimate before completion | Error 2 and error 5 . clicking noise when firing at high power .CTRL |  | BCTechSupport@iridex.com | 2025-02-13 13:09 | Yes | No |  | No |  | No
SRV-RQ002616 | WD | 65900-H500 | K1263 | CUST-LIO.WD, LOOSE MIRROR | Cannot test due to loose delivery mirror from flexure mirror support, loose beam | Replaced delivery mirror, ahlogen bulb, adjusted flexure mirror support, beam | BCTechSupport@iridex.com | 2025-02-13 13:01 | Yes | No |  | No |  | No
SRV-RQ002552 | PSD | 30903-H500 | 1358697 | CUST-LIO, PSD, ILLUMINATION CABLE BROKEN, | Cannot test due to Twisted and burnt assembly fiber optic, wrong and burnt |  | BCTechSupport@iridex.com | 2025-02-13 08:53 | Yes | No |  | No |  | No
SRV-RQ002662 | SD | 13030 | 29754-LS | CUST-SLx, PSD, VERY DIM OR NO AIMBEAM, URGENT REPAIR |  |  | BCTechSupport@iridex.com | 2025-02-13 08:15 | Yes | No |  | No |  | No
SRV-RQ002630 | PSD | 65511 | GH14022RM | CUST-SLA, PSD. (dropped) need a service estimate | BURNED FIBER JUMPER, LOOSE HANDLE HOLDER ON LEVER ASSY, |  | BCTechSupport@iridex.com | 2025-02-13 07:20 | Yes | No |  | No |  | No
SRV-RQ002569 | SD | 31808 | 110250F | CUST-FOOTSWITCH, PSD, DEAD , CALL FOR SHIP TO ADRESS |  |  | BCTechSupport@iridex.com | 2025-02-13 07:09 | Yes | No |  | No |  | No
SRV-RQ002540 | PSD | 30268-M | L2311 | CUST-LIO, LOW POWER, PSD | Cannot test due to Twisted and burnt assembly fiber optic, bent control stem, |  | BCTechSupport@iridex.com | 2025-02-12 09:35 | Yes | No |  | No |  | No
SRV-RQ001786 | PSD | 65515-H500 | 1380000312 | CUST-LIO. PSD. No beam. Fiber may be broken | Cannot test due to Twisted and broken assembly fiber optic, wrong and burnt |  | BCTechSupport@iridex.com | 2025-02-12 08:36 | Yes | No |  | No |  | No
SRV-RQ002354 | PSI | 66245-1 | IQI93430C | CUST-G6.PSI. Low power output | Power output low. SD | Realign and calibration Laser head. | BCTechSupport@iridex.com | 2025-02-12 06:26 | Yes | No |  | No |  | No
SRV-RQ002552 | PSD | 30903-H500 | 1358697 | CUST-LIO, PSD, ILLUMINATION CABLE BROKEN, | Cannot test due to Twisted and burnt assembly fiber optic, wrong and burnt |  | BCTechSupport@iridex.com | 2025-02-11 15:14 | Yes | No |  | No |  | No
SRV-RQ002550 | PSD | 87304 | 1101336819 | CUST-LIO, PSD, DOUBLE VISION CHECK ALIGNMENT | Cannot test due to Twisted assembly fiber optic, beam clipping, double vision no | Replaced assembly fiber optic, dust cap with lanyard, adjusted beam knob, | BCTechSupport@iridex.com | 2025-02-11 12:53 | Yes | No |  | No |  | No
SRV-RQ002573 | PSD | 13153-H | 1143484 | CUST-LIO,PSD, LOW POWER | Cannot test due to Twisted and broken assembly fiber, broken extension cable, |  | BCTechSupport@iridex.com | 2025-02-11 09:15 | Yes | No |  | No |  | No
SRV-RQ002353 | SI | 13030-220 | 30024-LS | CUST-SLX, FRONT LENSE DIRTY OR BURNT CHECK ALIGNMENT TOO, JUST REPAIRED RECENTLY |  |  | BCTechSupport@iridex.com | 2025-02-11 08:37 | Yes | No |  | No |  | No
SRV-RQ002541 | PSD | 30903-M | 1319902 | CUST-LIO, LOW POWER, PSD | Cannot test due to Twisted and burnt assembly fiber optic, extension cable white |  | BCTechSupport@iridex.com | 2025-02-10 15:11 | Yes | No |  | No |  | No
SRV-RQ002444 | PSI | 32000-4 | TX19159C | CUST-TX, PSI, Error 2 intermittent, (check laser head first pulse) | Can't dupplicated Error 2 . Just found Error 5,16 and error 38. ENT high dac . | Realign laser to manufacture TX specification. Yo | BCTechSupport@iridex.com | 2025-02-10 10:33 | Yes | No |  | No |  | No
SRV-RQ002264 | PSD | 30903-H500 | 1361749 | CUST-LIO,PSD, FRONT MIRROR BROKEN | Cannot test due to Twisted and burnt assembly fiber optic, burnt halogen bulb, |  | BCTechSupport@iridex.com | 2025-02-10 09:51 | Yes | No |  | No |  | No
SRV-RQ002658 | SD | 65600-01 | IQY71488C | CUST-IQ577, PSD, E05111 |  |  | BCTechSupport@iridex.com | 2025-02-10 08:13 | Yes | No |  | No |  | No
SRV-RQ002496 | SCD | 66245-1 | IQI93203C | CUST-CYCLO-G6, SCD, PM , VERIFY SHIP TO ADRESS BEFORE SGIPPING |  |  | BCTechSupport@iridex.com | 2025-02-07 08:03 | Yes | No |  | No |  | No
SRV-RQ002217 | PSI | 32000-1 | TX19235C | Cust-Tx, PSI, low power, provide estimate | Power shuts of when firing . Error 5,23,and 40 . Laser head require Eval. |  | BCTechSupport@iridex.com | 2025-02-07 07:25 | Yes | No |  | No |  | No
SRV-RQ002323 | PSD | 32000-1 | TX1293REVB | CUST-TX.PSD.LOW POWER (ERROR 2 | Error 2 and error 5 . Power keep reset when fring at high power ." message low | Realign laser head to manufacture TX specification. Yo | BCTechSupport@iridex.com | 2025-02-06 13:55 | Yes | No |  | No |  | No
SRV-RQ002528 | SI | 65600-08 | IQY72695C | Cust-IQ577, WI, LCM PCB and continious buzzering after turn on | NOT LOADING LASER MAY NEED LCM PCB HEAD SENT TO HEAD ROOM,JS | Replaced LCM 65105. Recalibration and tested power on the laser module. | BCTechSupport@iridex.com | 2025-02-06 12:32 | Yes | No |  | No |  | No
SRV-RQ002264 | PSD | 30903-H500 | 1361749 | CUST-LIO,PSD, FRONT MIRROR BROKEN | CANNOT TEST DUE BURNT FIBER AND BROKEN EXTENSION CBL. |  | BCTechSupport@iridex.com | 2025-02-06 09:45 | Yes | No |  | No |  | No
SRV-RQ002596 | SD | 65500 | IQG81586C | CUST-IQ532, PSD, UIM FAILED , JUST SERVICE AND SHIPPED | NOT BOOTING UP AS CUSTOMER COMPLAINT BECAUSE DAMAGED ENCODER |  | BCTechSupport@iridex.com | 2025-02-06 09:23 | Yes | No |  | No |  | No
SRV-RQ002467 | PSD | 70295 | YH14126TC | CUST-TxCell, PSD , CONECTIVITY ISSUE NOT RECOGNIZED, | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING |  | BCTechSupport@iridex.com | 2025-02-05 11:15 | Yes | No |  | No |  | No
SRV-RQ002623 | PSI | 66245-1 | IQI91980C | CUST-G6, E6002, PSI, Need estimate | E06006 Photocell detector readinge do not match | Replaced one optic (11586). Realign and calibration laser head. RY | BCTechSupport@iridex.com | 2025-02-05 11:04 | Yes | No |  | No |  | No
SRV-RQ002595 | PSD | 31174-01 | GH9335L | CUST-SLA, PSD, DAMAGED SMART KEY |  |  | BCTechSupport@iridex.com | 2025-02-04 13:39 | Yes | No |  | No |  | No
SRV-RQ002707 | SD | 65600-08 | IQY72518C | CUST-IQ577,PSD, POWER SURG HAPPENED NO DISPLAY | NOT BOOT UP BECAUSE SCM PCBA BLOWN UP | REPLACED SCM PCBA AND RETESTED ALL THE DELIVERY DEVICES.JS | BCTechSupport@iridex.com | 2025-03-03 07:29 | Yes | No |  | No |  | No
SRV-RQ002499 | SCD | 65900-H500 | 1352940 | CUST-LIO , SCD PP+, PM CHECKUP, CHECK SHIPPING ADDRESS | Cannot test due to Twisted and broken assembly fiber optic, broken control stem, |  | BCTechSupport@iridex.com | 2025-03-03 09:29 | Yes | No |  | No |  | No
SRV-RQ002722 | PSD | 65900-H500 | 1359450 | CUST-LIO, PSD, BROKEN KNOBS | Cannot test due to viewing separed from T-slot adapter, broken control stem, |  | BCTechSupport@iridex.com | 2025-03-03 10:29 | Yes | No |  | No |  | No
SRV-RQ002722 | PSD | 65900-H500 | 1359450 | CUST-LIO, PSD, BROKEN KNOBS | Cannot test due to viewing separate from T-slot adapter, broken control stem, |  | BCTechSupport@iridex.com | 2025-03-03 10:41 | Yes | No |  | No |  | No
SRV-RQ002704 | SD | 12733 | 2635-VET | CUST-DioVet, PSD, LPA | LOW POWER AS CUSTOMER COMPLAINT.HEAD SENT TOHEAD ROOM.JS |  | BCTechSupport@iridex.com | 2025-03-03 10:47 | Yes | No |  | No |  | No
SRV-RQ002754 | PSD | 65600-01 | IQY71616C | CUST-IQ577.psd. Communication problem. Check and estimate. | Communication problem per customer reported. | Technical Service requests to scrap the unit. DN | BCTechSupport@iridex.com | 2025-03-03 14:21 | Yes | No |  | No |  | No
SRV-RQ002751 | PSD | 30903-H | 1319902 | CUST-LIO, LOW POWER, PSD | Cannot test due to Twisted and burnt assembly fiber optic, extension cable white | Replaced assembly fiber optic, control stem, eccentric shaft, halogen bulb, | BCTechSupport@iridex.com | 2025-03-03 14:35 | Yes | No |  | No |  | No
SRV-RQ002659 | PSD | 31214 | GL15023F18 | CUST-GL, PSD, NO AIMBEAM POWERS OK, VERIFY SHIPPING ADRESS | Aiming Beam low power. 2.03 V, can't adjust Volt up 4.0 V . Laser head require |  | BCTechSupport@iridex.com | 2025-03-04 09:51 | Yes | No |  | No |  | No
SRV-RQ002417 | SD | 66245 | IQI91594C | GL, SL, G6 CEM,Calibration and PM |  |  | BCTechSupport@iridex.com | 2025-03-04 11:07 | Yes | No |  | No |  | No
SRV-RQ002743 | DD | 65777D | RC10274-D | DEMO REMOTE. DD. RETURN TO INVENTORY | Demo returned. SD | Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2025-03-04 12:52 | Yes | No |  | No |  | No
SRV-RQ002642 | PSD | 13030 | 30195-LS | CUST-SLX, PSD, No ILLUMINATION LIGHT | No ILLUMINATION LIGHT. SD | Replace power supply and retested all the delivery devices. | BCTechSupport@iridex.com | 2025-03-04 14:32 | Yes | No |  | No |  | No
SRV-RQ002681 | WD | 70292 | YEZ16317TC | CUST-TxCell, WD, E09514 x,y driver fault | CANNOT DUPLICATED THE PROBLEM. | TESTED ALIGNED CLEAN AND INSPECTED. | BCTechSupport@iridex.com | 2025-03-05 10:28 | Yes | No |  | No |  | No
SRV-RQ002699 | SD | 66245-1 | IQI92146C | CUST-CYCLOG6, PSD, UPON TURN ON IRIDEX LOGO PERSISTS |  |  | BCTechSupport@iridex.com | 2025-03-05 11:01 | Yes | No |  | No |  | No
SRV-RQ002757 | PSD | 31172-01 | GH9729LM | CUST-SLA.PSD.PM |  |  | BCTechSupport@iridex.com | 2025-03-06 12:25 | Yes | No |  | No |  | No
SRV-RQ002660 | SD | 66245-1 | IQI93526C | CUST-G6, PM( return ship see comments), PSD sticky knobs, please check front panel knob movements | UNABLE TO DUPLICATE STICKY KNOB BUT FOUND LOW POWER OUTPUT. RY | REALIGN AND CALIBRATION LASER HEAD. RY | BCTechSupport@iridex.com | 2025-03-06 13:41 | Yes | No |  | No |  | No
SRV-RQ002607 | PSD | 32000-4 | TX19422C | CUST-TX, PSD, KNOBS MISSING, NO SOUND, PLEASE EVALUATE | Misssing Duration and Interval knobs .#31783 . Retest and found volume control | Informed with Mohsin that the laser head has not completed eval. DN | BCTechSupport@iridex.com | 2025-03-07 10:12 | Yes | No |  | No |  | No
SRV-RQ002621 | STG | 65600-08 | IQY72634C | Cust IQ577 IQY72634C TxCell YEZ15309TC RC10328 STG, E05111 | EO5111 NOT DUPLICATED POWER WITH IN SPEC.JS | POWER WITH IN SPEC REMOTE AND FOOT SWITCH CAME WITH TESTED WORKS FINE.JS | BCTechSupport@iridex.com | 2025-03-07 12:00 | Yes | No |  | No |  | No
SRV-RQ002621 | STG | 65777 | RC10328 | Cust IQ577 IQY72634C TxCell YEZ15309TC RC10328 STG, E05111 | Remote came with console has no problem.js | Retested remote inhouse tool and with unit works fine.Js | BCTechSupport@iridex.com | 2025-03-07 12:01 | Yes | No |  | No |  | No
SRV-RQ000975 | PSI | 65600-01 | IQY71965C | Customer IQ577 PSI Laser Power Out Of Range error |  |  | BCTechSupport@iridex.com | 2025-03-07 13:47 | Yes | No |  | No |  | No
SRV-RQ002763 | PSI | 70292 | YEZ15408TC | CUST-TxCell, PSI, BROKEN FIBER | BROKEN FIBER THUMB CREW'S MISSING, BURNT FIBER | REPLACE FIBER OPTIC , REALIGNE THE TXCELL AND TESTED ACCORDING TO Mfg . SPEC | BCTechSupport@iridex.com | 2025-03-10 06:37 | Yes | No |  | No |  | No
SRV-RQ002510 | PSI | 66245-1 | IQI92751C | CUST-G6, PSI, PM ( software did not recognize Gprobe illuminate? per caller) | The complaint was not recognized Gprobe. Unable to duplicate the complaint. | Realign and calibration laser head and teated console with MFG equipment | BCTechSupport@iridex.com | 2025-03-10 07:10 | Yes | No |  | No |  | No
SRV-RQ002591 | SCD | 30903-H500 | 1319907 | CUST-LIO, SCD-GOLD, BROKEN KNOBS CAN NOT CONTROL IMAGE,VERIFY SHIP TO ADRESS | KLX and will need alignment. |  | BCTechSupport@iridex.com | 2025-03-10 16:11 | Yes | No |  | No |  | No
SRV-RQ002746 | PSD | 30903-H | 1170162 | CUSt-LIO.PSD. PM | Cannot test due to Twisted and burnt assembly fiber optic, burnt halogn bulb, |  | BCTechSupport@iridex.com | 2025-03-11 07:43 | Yes | No |  | No |  | No
SRV-RQ002597 | PSD | 65500 | IQG81443C | CUST-IQ532, PSD, NOT RECOGNIZING FIBERS | CUSTOMER COMPLAINT FIBER NOT RECOGNIZE BUT FOUND LOW POWER INSTEAD OF | -Realigned the laser head, recalibrated the unit, replaced fine adjust screws. | BCTechSupport@iridex.com | 2025-03-11 12:05 | Yes | No |  | No |  | No
SRV-RQ002761 | PSD | 30903-H500 | K1466 | cust-LIO.PSD.No ILLUMINATION LIGHT | No problem found with White light. | Replaced dusca with lanyard, adjusted flexure mirror and refocus laser spot size | BCTechSupport@iridex.com | 2025-03-11 12:20 | Yes | No |  | No |  | No
SRV-RQ002761 | PSD | 30903-H500 | K1466 | cust-LIO.PSD.No ILLUMINATION LIGHT | No problem found with White light. | Replaced dusca with lanyard, adjusted flexure mirror and refocus laser spot size | BCTechSupport@iridex.com | 2025-03-11 12:53 | Yes | No |  | No |  | No
SRV-RQ002766 | PSD | 11180-H500 | L2049 | CUST-LIO.PSD.PM | Cannot test due to Twisted and broken assembly fiber optic, bunt halogen bulb, |  | BCTechSupport@iridex.com | 2025-03-11 14:32 | Yes | No |  | No |  | No
SRV-RQ002647 | PSD | 13153-H | 1333608 | CUST-LIO, PSD, LOW POWER OUTPUT | Cannot test due to twisted and broken assembly fiber optic, burnt halogen bulb, |  | BCTechSupport@iridex.com | 2025-03-11 15:51 | Yes | No |  | No |  | No
SRV-RQ002784 | PSD | 65600-01 | IQY71881C | cust-IQ577.PSD TXCELL WAS HAVING ISSUE AND | Power output low |  | BCTechSupport@iridex.com | 2025-03-12 08:15 | Yes | No |  | No |  | No
SRV-RQ002570 | STG | 65600-01 | IQY72647C | CUST-IQ577, TxCell YEZ16332TC, RC10331, need to check together, after 11 shots cut off E05111 | Error E05111 per customer reported. | Replaced 65451. RETESTED ALL THE DELIVERY DEVICES AS MFG SPEC.JS | BCTechSupport@iridex.com | 2025-03-12 11:29 | Yes | No |  | No |  | No
SRV-RQ002648 | PSD | 66245-1 | IQI92571C | Q | Low power .laser head require Eval. CH.T | Realign and calibration laser head. RY | BCTechSupport@iridex.com | 2025-03-12 12:21 | Yes | No |  | No |  | No
SRV-RQ002168 | SCD | 30903-H500 | K1230 | CUST-LIO, SCD,PM CHECKUP | Cannot test due to extension cable damaged, Assembly fiber optic shows beam | Replaced assembly fiber optic, extension cable, adjusted flexure mirror support, | BCTechSupport@iridex.com | 2025-03-12 12:59 | Yes | No |  | No |  | No
SRV-RQ002570 | STG | 65777 | RC10331 | CUST-IQ577, TxCell YEZ16332TC, RC10331, need to check together, after 11 shots cut off E05111 | Remote came with console .js | Retested with console and inhouse tool as well.Js | BCTechSupport@iridex.com | 2025-03-12 13:00 | Yes | No |  | No |  | No
SRV-RQ002721 | SD | 65777 | RC0416 | CUST-REMOTE, PSD, BROKEN DROP TREAT to STD. BY PLEASE CK W/ CUST IQY71352C | Remote came with console and has speaker became loose. | Reinstalled to place and secured to avoid become loose and retested .JS | BCTechSupport@iridex.com | 2025-03-12 13:05 | Yes | No |  | No |  | No
SRV-RQ002691 | SD | 66245-1 | IQI93219C | CUST-CYCLO-G6, PM, PSD. VERYFY SHIP TO BEFORE SHIPPING | PM CHECK.JS | RETESTED ALL THE DELIVERY DEVICES .JS | BCTechSupport@iridex.com | 2025-03-13 07:19 | Yes | No |  | No |  | No
SRV-RQ002666 | SD | 65500-08 | IQG81890C | CUST-IQ532, PSD, FIBER FERULE CHECK FIBER WON'T STAY ON ITS PLACE | Encoder duration knob was not working during process found uim board has damage |  | BCTechSupport@iridex.com | 2025-03-13 09:46 | Yes | No |  | No |  | No
SRV-RQ002772 | SD | 12731 | 2694 | CUST-SL.CEM.CALIBRATION AND PM |  |  | BCTechSupport@iridex.com | 2025-03-14 07:37 | Yes | No |  | No |  | No
SRV-RQ002668 | PSD | 66245-1 | IQI92869C | CUST-CYCLO-G6, PSD, PM CHECKUP | Power output low | Realign and calibration laser head. RY | BCTechSupport@iridex.com | 2025-03-14 08:30 | Yes | No |  | No |  | No
SRV-RQ002810 | PSD | 30903-H500 | 1361063 | CUST-LIO.PSD.POWER OUTPUT WAS LOW MUST CHECK IF FIBER COMPROMISED | Cannot test due to Twisted and burnt assembly fiebr optic, bent control stem, |  | BCTechSupport@iridex.com | 2025-03-14 10:53 | Yes | No |  | No |  | No
SRV-RQ002611 | STG | 65777 | RC10285 | Cust remotes (3) need PM and testing. Upgrade to the Latest IQ compatibility SW | Customer return for upgrade software .sd | Upgrade to 2.2.1 software . SD | BCTechSupport@iridex.com | 2025-03-17 08:24 | Yes | No |  | No |  | No
SRV-RQ002611 | STG | 65777D | RC0338D | Cust remotes (3) need PM and testing. Upgrade to the Latest IQ compatibility SW | Need upgrade software require by customer. jumpy encoder.SD | Upgrade 2.2.1 software and replace encoder PCAB .SD | BCTechSupport@iridex.com | 2025-03-17 14:33 | Yes | No |  | No |  | No
SRV-RQ002798 | PSD | 30903-H500 | 1380000360 | CUST-LIO, PSD, ILLUM. CABLE DISCONNECTED | Cannot test due to Twisted and burnt assebly fiber optic, damaged extension |  | BCTechSupport@iridex.com | 2025-03-18 14:02 | Yes | No |  | No |  | No
SRV-RQ002695 | PSD | 65900-H500 | K1794 | CUST-LIO, PSD, LIGHT IS TOO BRIGHT | Cannot test due to Missing eye safety filters 577nm, +2 dioter lens, burnt |  | BCTechSupport@iridex.com | 2025-03-18 12:08 | Yes | No |  | No |  | No
SRV-RQ002321 | WI | 32000-4 | TX19560C | cust-TX.WI."Low Power" | Low power and high dac on ENT . Laser head require Eval. CT | Realign laser head to manufacture Tx specification. Yo | BCTechSupport@iridex.com | 2025-03-18 10:45 | Yes | No |  | No |  | No
SRV-RQ002783 | SD | 65500 | IQG81289C | CUST-IQ532, PSD, NOT RECOCNIZING LIO & SLA | Duplicated customer complaint intermittengly recognise fiber. |  | BCTechSupport@iridex.com | 2025-03-18 09:39 | Yes | No |  | No |  | No
SRV-RQ002768 | SD | 13030 | 29453-LS | CUST- It was experiencing power loss during a case and needs evaluation/repair. | CUSTOMER COMPLAINT LOW POWER AND SEEN ERROR 5 IN STORAGE BUT | Recalibrated all the delivery devices and photocell and retested . Foot switch | BCTechSupport@iridex.com | 2025-03-18 05:53 | Yes | No |  | No |  | No
SRV-RQ002762 | SI | 66245-1 | IQI92176C | CUST- G6. . WI, PM, may have May have SCM battery problem, please check and repair | Dead battery E05103.Need to replace battery. After the battery replaced same | Replaced scm pcba and retested all the delivery devices JS | BCTechSupport@iridex.com | 2025-03-19 08:19 | Yes | No |  | No |  | No
SRV-RQ002775 | SD | 66245-1 | IQI91945C | CUST-CYCLO-G6, SCD, PM , VERYFY SHIP TO ADRESS | PM CHECK.JS | Retested all the delivery devices as MFG spec. JS | BCTechSupport@iridex.com | 2025-03-19 11:02 | Yes | No |  | No |  | No
SRV-RQ002669 | PSD | 30903-H500 | K2104 | CUST-LIO, PSD, DOUBLE VISION | Cannot test due to Loose housing illumination and missing phillip screw, loose |  | BCTechSupport@iridex.com | 2025-03-19 15:47 | Yes | No |  | No |  | No
SRV-RQ002669 | PSD | 30903-H500 | K2104 | CUST-LIO, PSD, DOUBLE VISION | Cannot test due to Loose housing illumination and missing phillip screw, loose | Replaced assembly fiber optic, halogen bulb,dust cap with lanayard, reglued | BCTechSupport@iridex.com | 2025-03-19 15:58 | Yes | No |  | No |  | No
SRV-RQ002619 | PSI | 66120 | IQHG81033C | CUST-IQ532.PSD.PM | -E06006 photocell mismatch error | -Realigned the laser module, recalibrated the unit. Verified all the functions, | BCTechSupport@iridex.com | 2025-03-20 08:29 | Yes | No |  | No |  | No
SRV-RQ002609 | WI | 65600-08 | IQY72767C | CUST-IQ577 and YEZ16627TC ( single spot aiming beam issue ) + loaner return TxCell | Low aiming beam | -Realigned the laser module, recalibrated the unit. Verified all the functions | BCTechSupport@iridex.com | 2025-03-21 07:09 | Yes | No |  | No |  | No
SRV-RQ002696 | SD | 65600-08 | IQY72631C | CUST-IQ577, PSD, BEEPING , VERIFY SHIPPING ADRESS BEFOR SHIPPING | Display beeping need to replace display and gasket.JS | Realign and calibration port 1 and port 2. | BCTechSupport@iridex.com | 2025-03-21 07:11 | Yes | No |  | No |  | No
SRV-RQ002715 | SI | 13030 | 30257-LS | Cust-SLx, LIO ( S/N ?), PSI, PM and LIO LUX needs to be checked for >2.7 Klux | PM CHECK AND LUX LIGHT CHECK.JS | Recalibrated all the delivery devices and checked lio light as MFG spec. JS | BCTechSupport@iridex.com | 2025-03-21 07:40 | Yes | No |  | No |  | No
SRV-RQ002828 | PSD | 66245-1 | IQI93228C | Cust-G6 console (PM $2800.00) | Customer return for PM check .SD | Retested all the delivery devices.SD | BCTechSupport@iridex.com | 2025-03-21 12:02 | Yes | No |  | No |  | No
SRV-RQ002811 | SD | 65500 | IQG81088C | CUST-IQ532.PSD. SCREEN WENT BLANK WHEN PLUG PROBE IN | Screen went blank customer complaint Have not duplicated customer complaint. JS | Retested all the delivery devices as power with in MFG spec. JS | BCTechSupport@iridex.com | 2025-03-21 12:30 | Yes | No |  | No |  | No
SRV-RQ002845 | PSD | 30268-H | L2311 | CUST-LIO, LOW POWER, PSD | Cannot test due to Twisted and burnt assembly fiber optic, bent control stem, | Replaced assembly fiber optic, halogen bulb, control stem, ecentric shaft, Heine | BCTechSupport@iridex.com | 2025-03-21 13:06 | Yes | No |  | No |  | No
SRV-RQ002857 | DD | 66245D | IQI91587C-D | DEMO-G6.DD.RETURN INTO INVENTORY |  |  | BCTechSupport@iridex.com | 2025-03-21 13:53 | Yes | No |  | No |  | No
SRV-RQ002756 | SD | 12733 | 2605-VET | CUST-VET, LPA, PSD, | NO AIMING BEAM AND LOW POWER LOOSE OPTIC IN BOX,JS | Reinstall aiming beam lens holder and realign laser head. RY | BCTechSupport@iridex.com | 2025-03-27 13:36 | Yes | No |  | No |  | No
SRV-RQ002827 | SD | 65500 | IQG81079C | CUST-IQ532,PSD, JUMPY ENCODER | Bad encoder as customer complaint.and low power as well head sent to head room J | -Replaced fine adjust screws, realigned the laser head, recalibrated the unit. | BCTechSupport@iridex.com | 2025-03-27 10:22 | Yes | No |  | No |  | No
SRV-RQ002701 | PSD | 70292 | YEZ14137TC | CUST-TxCell, PSD, E05111 | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING | REPLACED COMPUTER BOARD BEZELUSB CABLE TESTED ALIGNED CALIBRATED | BCTechSupport@iridex.com | 2025-03-27 07:04 | Yes | No |  | No |  | No
SRV-RQ002478 | PSI | 66245-1 | IQI92424C | CUST-CYCLO-G6, PSI , E05103 DEAD BATTERY MESSAGE | E05103 DEAD BATTERY MESSAGE |  | BCTechSupport@iridex.com | 2025-03-26 14:09 | Yes | No |  | No |  | No
SRV-RQ002731 | PSD | 31172-01 | GH11208LM | CUST-TX, PSD, NO AIMBEAM CK WITH SLA RMA02467 | No Aimbeam CK. PM SLA, Clean Mirrors, Lenses and Test. | No Problem was found. compleated PM on SLA, leaned Mirors, Lenses and | BCTechSupport@iridex.com | 2025-03-26 12:13 | Yes | No |  | No |  | No
SRV-RQ002834 | PSD | 30903-H500 | K1095 | CUST-LIO, PSD, BROKEN ILLUMINATION CABLE | Cannot test due to broken extension cable, loose aiming beam knob adjust, burnt |  | BCTechSupport@iridex.com | 2025-03-26 11:32 | Yes | No |  | No |  | No
SRV-RQ002570 | STG | 70292 | YEZ16332TC | CUST-IQ577, TxCell YEZ16332TC, RC10331, need to check together, after 11 shots cut off E05111 | BURNED FIBER. | REPLACED FIBER TESTED ALIGNED CALIBRATED CLEAN AND INSPECTED. | BCTechSupport@iridex.com | 2025-03-26 10:46 | Yes | No |  | No |  | No
SRV-RQ002797 | PSD | 30903-H500 | 1358694 | CUST-LIO, PSD, BROKEN KNOBS, VERIFY SHIP TO ADRESS | Cannot test due to Twisted assembly fiber optic, broken control stem, broken |  | BCTechSupport@iridex.com | 2025-03-26 08:33 | Yes | No |  | No |  | No
SRV-RQ002801 | PSD | 65500-08 | IQG81967C | CUST-IQ532, PSD, EO511, PORT 1 NOT RECOGNIZING, VERIFY SHIP TO | EO511, PORT 1 NOT RECOGNIZING. SD | No problem found on EO511, PORT 1 NOT RECOGNIZING. | BCTechSupport@iridex.com | 2025-03-26 08:29 | Yes | No |  | No |  | No
SRV-RQ002840 | PSD | 30903-H | L2873 | Cust-LIO, PSD, loose item inside the LIO, Provide an estimate | Cannot test due to Twisted and burnt assembly fiber optic, broken assembly fiber |  | BCTechSupport@iridex.com | 2025-03-25 15:49 | Yes | No |  | No |  | No
SRV-RQ002787 | PSD | 30903-H500 | 1347884 | CUST-LIO, PSD, LOW POER OUT PUT, VERIFY SHIP TO | Cannot test due to Twisted and broken assembly fiber optic, bent control stem, |  | BCTechSupport@iridex.com | 2025-03-25 14:06 | Yes | No |  | No |  | No
SRV-RQ002547 | PSI | 70292 | YEZ14974TC | CUST-TxCell, PSI,LED BLINKING INTERMITTENTLY, sCANNER BOARD REPLACED | CONTROL NON-STOP BLINKING BAD USB CABLE. | REPLACED USB CABLE'S TESTED ALIGNED CALIBRATED CLEAN AND INSPECTED. | BCTechSupport@iridex.com | 2025-03-25 12:31 | Yes | No |  | No |  | No
SRV-RQ002837 | PSD | 30903-H | 1264712 | CUST-LIO.PSD.PM (Incompatible Device ISSUE WITH GLX?) | Cannot test due to twisted and burnt aseembly fiber optic, bent control stem, |  | BCTechSupport@iridex.com | 2025-03-25 11:07 | Yes | No |  | No |  | No
SRV-RQ002837 | PSD | 30903-H | 1264712 | CUST-LIO.PSD.PM (Incompatible Device ISSUE WITH GLX?) | Cannot test due to twisted and burnt aseembly fiber optic, bent control stem, |  | BCTechSupport@iridex.com | 2025-03-25 11:03 | Yes | No |  | No |  | No
SRV-RQ002808 | SD | 65600-01 | IQY71683C | CUST-IQ577, PSD, POWER INLET BROKEN | Customer complaint broken power inlet.Js | Have not duplicated any problem inlel intact with firmly to back panel | BCTechSupport@iridex.com | 2025-03-25 10:59 | Yes | No |  | No |  | No
SRV-RQ002827 | SD | 65500 | IQG81079C | CUST-IQ532,PSD, JUMPY ENCODER |  |  | BCTechSupport@iridex.com | 2025-03-25 07:07 | Yes | No |  | No |  | No
SRV-RQ002667 | SD | 13030 | 27371 | CUST-SLX, PSD, PM CHECKU |  |  | BCTechSupport@iridex.com | 2025-03-25 06:53 | Yes | No |  | No |  | No
SRV-RQ002599 | PSD | 31808 | 112644F | CUST-FOOTSWITCH, NO COMMUNICATION, | Red light Flash blinking . No communication .Send back to suppiler |  | BCTechSupport@iridex.com | 2025-03-25 06:27 | Yes | No |  | No |  | No
SRV-RQ002851 | STG | 65600-08 | IQY72578C | Cust-TxCell, IQ, TxCell does not boot, STG, remote control only briefly displays the IRIDX symbol |  |  | BCTechSupport@iridex.com | 2025-03-24 12:19 | Yes | No |  | No |  | No
SRV-RQ002686 | PSD | 66245-1 | IQI92877C | CUST-CYCLO-G6, PSD, Dropping Treat to Stdby, confirm shipping adress before shipping | Dropping Treat to Stdby. SD | No problem found. Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2025-03-24 11:46 | Yes | No |  | No |  | No
SRV-RQ002888 | PSD | 12729 | 27371 | CUST-SLX, PSD, PM CHECKU | Low power output. | Replaced assy. L1 lens, adjust screws and realign laser head. RY | BCTechSupport@iridex.com | 2025-03-31 10:55 | Yes | No |  | No |  | No
SRV-RQ002852 | PSD |  |  | CUST-SL, PSD, PM CHECKUP |  |  | BCTechSupport@iridex.com | 2025-03-31 14:27 | Yes | No |  | No |  | No
SRV-RQ002529 | PSD | 87304 | 1101336801 | CUST-LIO, PSD, MIRROR LOOSE INSIDE | Cannot test due to burnt aseenbly fiber optic, broken boot, loose mirror from |  | BCTechSupport@iridex.com | 2025-04-01 09:41 | Yes | No |  | No |  | No
SRV-RQ002501 | PSI | 70297 | EZZ13156TC | CUST-TxCell, PSI, E05111 AND EVALUATE | CONTROL BOX NON-STOP BLINKING |  | BCTechSupport@iridex.com | 2025-04-01 11:08 | Yes | No |  | No |  | No
SRV-RQ002665 | STG | 70292 | YEZ16147TC | Cust-TxCell, IQ, TxCell does not boot, STG, remote control only briefly displays the Iridex symbo | COMPUTER BOARD NOT WORLINK PROPERLY. | REPLACED COMPUTER BOARD TESTED ALIGNED CALIBRATED CLEAN AND INSPECTED. | BCTechSupport@iridex.com | 2025-04-01 13:18 | Yes | No |  | No |  | No
SRV-RQ002639 | WD | 13152-H500 | K2442 | CUST-LIO, broken/damaged fiberoptic/PSD |  |  | BCTechSupport@iridex.com | 2025-04-01 14:41 | Yes | No |  | No |  | No
SRV-RQ002759 | PSD | 70292 | YEZ14609TC | CUST-TxCell, PSD, DIM AIM BEAM & UNKNOWN ERROR | UNABLE TO TESTED BURNED FIBER, UPGRADED INTERFACE BOARD AND USB CABLE, |  | BCTechSupport@iridex.com | 2025-04-02 06:29 | Yes | No |  | No |  | No
SRV-RQ002711 | PSD | 31215 | GL15210C30 | CUST-GLx , PSD, PM CHECKUP, verify shipp to before shipping | Low power on 2000mw . | Realign laser head to manufacture GL 3rd cooler specification. Yo | BCTechSupport@iridex.com | 2025-04-02 06:56 | Yes | No |  | No |  | No
SRV-RQ002793 | PSD | 70295 | YH14841TC | cust-IQ577.PSD TXCELL WAS HAVING ISSUE AND | NO PROBLEM FOUND. | TESTED CALIBRATED CLEAN AND INSPECTED. | BCTechSupport@iridex.com | 2025-04-02 07:58 | Yes | No |  | No |  | No
SRV-RQ002843 | PSD | 30903-H500 | K0926 | Cust-LIO, PSD, (S/N does not bring recdords) | Cannot test due to Burnt assembly fiber optic, missing eyepiece rubber, burnt |  | BCTechSupport@iridex.com | 2025-04-02 10:35 | Yes | No |  | No |  | No
SRV-RQ002843 | PSD | 30903-H500 | K0926 | Cust-LIO, PSD, (S/N does not bring recdords) | Cannot test due to Burnt assembly fiber optic, missing eyepiece rubber, burnt |  | BCTechSupport@iridex.com | 2025-04-02 11:00 | Yes | No |  | No |  | No
SRV-RQ002859 | PSI | 13152-H500 | K1397 | CUST-LIO.PSI. LUX needs to be checked for >2.7 Klux | No problem found with white light, White light read 2.48 KLx spec is 2.00 KLx | Replaced halogen bulb, dust cap with lanyar, adjusted flexure mirror support, | BCTechSupport@iridex.com | 2025-04-02 13:52 | Yes | No |  | No |  | No
SRV-RQ002855 | SD |  |  | CUST-LIO.PSD. PM | adjustement, realignment and refocus. |  | BCTechSupport@iridex.com | 2025-04-02 13:54 | Yes | No |  | No |  | No
SRV-RQ002855 | SD |  |  | CUST-LIO.PSD. PM | adjustement, realignment and refocus. |  | BCTechSupport@iridex.com | 2025-04-02 13:54 | Yes | No |  | No |  | No
SRV-RQ002855 | SD | 11180-H | L1779 | CUST-LIO.PSD. PM | adjustement, realignment and refocus. |  | BCTechSupport@iridex.com | 2025-04-02 13:54 | Yes | No |  | No |  | No
SRV-RQ002790 | PSD | 30903-H500 | K2112 | Cust-LIO, PSD, missing ESF or a lens, provide estimate before repair | Cannot test due to damaged extension cable, missing eye safety filter 532nm, |  | BCTechSupport@iridex.com | 2025-04-02 14:42 | Yes | No |  | No |  | No
SRV-RQ002820 | PSD | 70300 | GH14540TC | CUST-TxCell, PSD, LED BLINKING | CONTROL BOX NON-STOP BLINKING,AND ERROR E09505 BAD GALVO'S |  | BCTechSupport@iridex.com | 2025-04-03 07:44 | Yes | No |  | No |  | No
SRV-RQ002726 | PSD | 31214 | GL16053F30 | CUST-GL 2007 model, PSD, PM ( came with SLA GH5120, RMA02291) PSD | LIO hight dac out of spect. (0.97mw) . and power are not stable . | Realign laser head to manufacture GL 3rd cooler specification. Yo | BCTechSupport@iridex.com | 2025-04-03 08:05 | Yes | No |  | No |  | No
SRV-RQ002864 | WD | 13030 | 30447-LS | CUST-SLx, WD, NO WHITE LIGHT & KNOBS NOT WOKING |  |  | BCTechSupport@iridex.com | 2025-04-03 10:15 | Yes | No |  | No |  | No
SRV-RQ002825 | WI | 70297 | EZZ16441TC | CUST-TxCell, WI, NEED EVALUATION | UNABLE TO TESTED ASSY,CABLE CIRCULAR CONN IS NOT WORKING. |  | BCTechSupport@iridex.com | 2025-04-03 12:34 | Yes | No |  | No |  | No
SRV-RQ002889 | WD | 65511 | GH16850RM | CUST-SLA.WD.BROKEN LENS | ESF filter is broken. Align and Calibrate. Clean mirrors, lenses and Test. |  | BCTechSupport@iridex.com | 2025-04-03 12:46 | Yes | No |  | No |  | No
SRV-RQ002568 | WD | 32000-4 | TX19598C | CUST-TX, WD, PM AND NEED SERVICE, VERYFY SHIPPIN ADRESS BEFORE SHIPPING | Stripped Ferrule ,can not operation . Laser head require Eval. CT | The unit is under warranty. Please move forward with your repair. DN | BCTechSupport@iridex.com | 2025-04-04 07:41 | Yes | No |  | No |  | No
SRV-RQ002724 | PSD | 70292 | YEZ14322TC | CUST-TxCell, PSD, DOES NOT RECOGNIZED MAY BE FIBER ISSUE | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING COMPUTER BAORD NOT WORKING |  | BCTechSupport@iridex.com | 2025-04-04 08:33 | Yes | No |  | No |  | No
SRV-RQ002786 | WD | 65600D-01 | IQY71712C-D | CUST-IQ577.SCD.WON'T STAY ON HAD TO HOLD KEY |  |  | BCTechSupport@iridex.com | 2025-04-08 07:59 | Yes | No |  | No |  | No
SRV-RQ002835 | PSD | 31808 | 112747F | CUST-FOOTSWITCH, NEED BATTERIES | Red light blinking , low batterry . require send back to suplier for Eval. CT | Send back to supplier for repair . CT | BCTechSupport@iridex.com | 2025-04-08 08:47 | Yes | No |  | No |  | No
SRV-RQ002920 | SD | 66245 | IQI91124C | CUST-G6.PSD.UNIT NOT ACCEPTING PROBE | Not duplicated customer complaint fiber works fine.JS | Retested all the delivery devices as MFG spec.JS | BCTechSupport@iridex.com | 2025-04-08 09:00 | Yes | No |  | No |  | No
SRV-RQ002926 | PSD | 31214 | GL15023F18 | CUST-GL, PSD, NO AIMBEAM POWERS OK, VERIFY SHIPPING ADRESS | Aiming Beam low power. 2.03 V, can't adjust Volt up 6.5 V. | Realign laser head to manufacture GL 3rd cooler specification. Yo | BCTechSupport@iridex.com | 2025-04-08 10:07 | Yes | No |  | No |  | No
SRV-RQ002887 | PSD | 70295 | YH14234TC | CUST-TxCell, PSD, AUX DEVICE,, led BLINKING, VERIFY ADRESS BEFORE SHIPPING | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING |  | BCTechSupport@iridex.com | 2025-04-08 13:01 | Yes | No |  | No |  | No
SRV-RQ002689 | PSD | 12731 | 1461 | CUST-SL, PSD, PM CHECKUP | Customer return for PM check. SD | Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2025-04-08 13:53 | Yes | No |  | No |  | No
SRV-RQ002787 | PSD | 30903-H500 | 1347884 | CUST-LIO, PSD, LOW POER OUT PUT, VERIFY SHIP TO | Cannot test due to Twisted and broken assembly fiber optic, bent control stem, |  | BCTechSupport@iridex.com | 2025-04-08 15:02 | Yes | No |  | No |  | No
SRV-RQ002673 | WI | 65600-08 | IQY72623C | CUST-IQ577 Topcon IQY72623C, bad port one, WI | Bad port one and power output low | Replace resistance fiber detect board. SD | BCTechSupport@iridex.com | 2025-04-09 10:11 | Yes | No |  | No |  | No
SRV-RQ002624 | PSD | 66245-1 | IQI92094C | Cust-G6, PM, PSD, Urgent PM needed, $2800 and if problems found, provide estimate. | Customer return for PM check. SD | Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2025-04-10 10:53 | Yes | No |  | No |  | No
SRV-RQ002813 | PSD | 30903-H | L2956 | CUST- OPTHALMOSCOPE, PSD, LASER AIMING BEAM WOULD NOT ADJUST TO THE CENTER | Cannot test due to Twisted and broken assembly fiber optic, burnt halogen bulb, |  | BCTechSupport@iridex.com | 2025-04-10 14:25 | Yes | No |  | No |  | No
SRV-RQ002567 | WD | 32000-4 | TX19594C | CUST-TX, WD, PM AND NEED SERVICE, VERYFY SHIPPIN ADRESS BEFORE SHIPPING | Power low at 2500mw .Laser head require Eval. CT | The unit is under warranty, please move forward with your repair. DN | BCTechSupport@iridex.com | 2025-04-10 14:43 | Yes | No |  | No |  | No
SRV-RQ002792 | PSD | 31214 | GL13766F | CUST-GL, PSD, FIBER FOCUS FERRULE STRIP | Stripped Ferrule can't firing .Laser head require Eval. CT | Realign laser head to manufacture GL 3rd cooler specification. Yo | BCTechSupport@iridex.com | 2025-04-11 14:12 | Yes | No |  | No |  | No
SRV-RQ002180 | WD | SA-06568 | 90012924 | IRIDEX FGI.SA-06568 IP 577. WD/Sales Return for Investigation, No Power To Sys/Rework/Return to FGI | No power reported per customer. | Replaced laser diode,TEC and insulation. TN | BCTechSupport@iridex.com | 2025-04-14 07:18 | Yes | No |  | No |  | No
SRV-RQ002805 | PSD | 30905D | GL14757F30-D | CUST-GL. PSD. Stripped PORT? | Stripped Ferrule. Laser head require Eval . CT | Realign laser head to manufacture GL 3rd cooler specifcation. Yo | BCTechSupport@iridex.com | 2025-04-14 07:30 | Yes | No |  | No |  | No
SRV-RQ002799 | SD | 13030 | 27444-LS | CUST-SLX. PSD. LIO LIGHT CONNECTOR IS LOOSE | No illumination light and loose RCA jack.js | Re-soldered rca jack and retighten, recalibrated all the delivery devices. JS | BCTechSupport@iridex.com | 2025-04-14 09:36 | Yes | No |  | No |  | No
SRV-RQ002838 | PSD | 65600-01 | IQY71518C | CUST-IQ577,PSD,DUE TO POWER SURGE NO BOOTUP OT DISPLAY | NO BOOTUP OT DISPLAY.SD | No problem found on LCD display. Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2025-04-14 10:50 | Yes | No |  | No |  | No
SRV-RQ002880 | SD | 13030 | 30385-LS | CUST-SLX.WD.Unable to get the LED light bulb to illuminate on the headset | NO illumination light as customer complaint.js | Replaced power supply and recalibrated all the delivery devices as per MFG spec, | BCTechSupport@iridex.com | 2025-04-14 12:54 | Yes | No |  | No |  | No
SRV-RQ002945 | PSI | 32002 | TX18122C | CUST-HEAD,PSD, LPA300MMW | Vanadate crack and fiber adapter stripped. Need to replace them both. YO | Cust. approved. DN | BCTechSupport@iridex.com | 2025-04-14 13:51 | Yes | No |  | No |  | No
SRV-RQ002502 | PSD | 66245-1 | IQI93132C | cust-G6.PSD.PM | POWER LOW AT jumper 100um. Laser head require Eval CT | Replaced cylinder lens assy. and fiber adpater and realign and calibration. | BCTechSupport@iridex.com | 2025-04-15 08:08 | Yes | No |  | No |  | No
SRV-RQ002961 | LD | 65511-L | GH13426RM-L | LOANER-SLA.LD.RETURN INTO INVENTORY |  |  | BCTechSupport@iridex.com | 2025-04-15 10:19 | Yes | No |  | No |  | No
SRV-RQ002826 | WI | 65500-08 | IQG82179C | CUST-IQ532. WI. the display beeps erratically ( request for WR ) |  |  | BCTechSupport@iridex.com | 2025-04-15 12:26 | Yes | No |  | No |  | No
SRV-RQ002814 | SD | 31214 | GL14059F30 | CUST-GL, PSD, READING POWERS TOO HIGH, PLEASE CALIBRATE IN THE MIDLE OF RANG | High power per customer reported. CT | Realign laser head to manufacture GL 3rd cooler specification. Yo | BCTechSupport@iridex.com | 2025-04-15 12:57 | Yes | No |  | No |  | No
SRV-RQ002912 | SD | 13030 | 29424-LS | CUST-SLx, PSD, enters standby after first pulse. | After one pulse goes to stand by,because of no calibration data in storage may | Replaced bench mark and recalibrated all the delivery devices.JS | BCTechSupport@iridex.com | 2025-04-15 13:45 | Yes | No |  | No |  | No
SRV-RQ002954 | SD | 66245D-1 | IQI92520C-D | DEMO-G6.DD.RETURN INTO INVENTORY | Demo return.SD | Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2025-04-16 07:19 | Yes | No |  | No |  | No
SRV-RQ002947 | PSD | 66245 | IQI91166C | CUST-CYCLO-G6,PSD, PM AND GUI UPGRADE | Customer return for upgrade soft ware. SD | Upgrade to quail 3 soft ware. SD | BCTechSupport@iridex.com | 2025-04-16 09:50 | Yes | No |  | No |  | No
SRV-RQ002682 | SI | 65500 | IQG81000C | Cust-IQG532, PSI, Mfg. 2010, Low power error, need estimate or indicate can not repair | As customer complaint low power E06002 Appears as soon as pressed foot switch.HR |  | BCTechSupport@iridex.com | 2025-04-16 15:14 | Yes | No |  | No |  | No
SRV-RQ002608 | PSD | 70297 | EZZ14546TC | CUST-TxCell,PSD, E05111 | CONTROL BOX NON-STOP BLINKING ERROR (E05111) BAD COMPUTER BOARD |  | BCTechSupport@iridex.com | 2025-04-17 07:58 | Yes | No |  | No |  | No
SRV-RQ002969 | PSI | 65511 | GH11100RM | CUST-SLA.PSI.PM |  |  | BCTechSupport@iridex.com | 2025-04-17 09:56 | Yes | No |  | No |  | No
SRV-RQ002883 | SD | 66245-1 | IQI92134C | CUST-G6, PSD, UIM is not loading | UIM FAIL TO LOAD APPLICATION AS CUSTOMER COMPLAINT.JS | Replaced fiber adpater,lens molded realign and calibration. RY | BCTechSupport@iridex.com | 2025-04-17 12:16 | Yes | No |  | No |  | No
SRV-RQ002758 | PSD | 31214 | GL15736F18 | CUST-GL.PSD.PM | ERROR 5 low power and intermittent heating delta . found heater board 30791 bad. | Realign laser head to manufacture GL 3rd cooler specification. Yo | BCTechSupport@iridex.com | 2025-04-17 12:42 | Yes | No |  | No |  | No
SRV-RQ002910 | PSD | 66245 | IQI91301C | CUST-G6.PSD PM CHECKUP | POWER OUTPUT LOW | Realign and calibration. RY | BCTechSupport@iridex.com | 2025-04-18 06:07 | Yes | No |  | No |  | No
SRV-RQ002714 | PSI | 12732 | 1424-VET | CUST-DIOVET. PSI. Low power, S/N# DioVet1424, Contact: Naina | Low power output and broken lens display. SD | Recalibration all the output and replace broken lens display .SD | BCTechSupport@iridex.com | 2025-04-18 07:00 | Yes | No |  | No |  | No
SRV-RQ002786 | WD | 65600D-01 | IQY71712C-D | CUST-IQ577.SCD.WON'T STAY ON HAD TO HOLD KEY | WON'T STAY ON HAD TO HOLD KEY and POWER OUTPUT LOW.SD |  | BCTechSupport@iridex.com | 2025-04-18 08:45 | Yes | No |  | No |  | No
SRV-RQ002675 | WI | CUST-IQ577 | IQY72787C | CUST-IQY72787 , fan speed high, WI | -fan running high | The LCM board doesn't function properly. | BCTechSupport@iridex.com | 2025-04-18 11:57 | Yes | No |  | No |  | No
SRV-RQ002895 | WD | 66245-1 | IQI92196C | CUST-G6, Need GUI upgrade, WD ( probe deal?!). |  |  | BCTechSupport@iridex.com | 2025-04-18 13:03 | Yes | No |  | No |  | No
SRV-RQ002968 | SD | 13030 | 29693-LS | CUST-SLX.CEM.Calibration and PM | Annual PM. RY | Reteste per MFG test traveler spce. RY | BCTechSupport@iridex.com | 2025-04-21 08:09 | Yes | No |  | No |  | No
SRV-RQ002967 | SD | 13030D | 25759-D | CUST-SL.CEM.CALIBRATION | PM CHECK | RETESTED AS MFG SPEC.JS | BCTechSupport@iridex.com | 2025-04-21 08:26 | Yes | No |  | No |  | No
SRV-RQ002919 | PSD | CUST-TXCELL | GH14254TC | CUST, PSD, SLIT LAMP ATTACHMENT AND CONNECTING BOX STOPPED WORKING | UNABLE TO TESTED COMPTROL BOX NON-STOP BLINKING |  | BCTechSupport@iridex.com | 2025-04-21 10:59 | Yes | No |  | No |  | No
SRV-RQ002718 | SI | CUST-CYCLO-G6 | IQI93771C | CUST-G6.PSI.PROBES NOT BEING RECOGNIZE BY CONSOLE | RFID FIFER TESTED WORKS FINE NO PROBLEM FOUND.JS | RETESTED ALL REQUIRE DELIVERY DEVICES AS MFG SPEC.JS | BCTechSupport@iridex.com | 2025-04-21 11:37 | Yes | No |  | No |  | No
SRV-RQ002730 | PSD | 32000-1 | TX1440REVB | CUST-TX, PSD, NO AIMBEAM CK WITH SLA RMA02468 | Low Aiming Beam on SLA .Laser head require Eval. ct | Realigned laser head to manufacture TX spec. YO | BCTechSupport@iridex.com | 2025-04-22 05:30 | Yes | No |  | No |  | No
SRV-RQ002765 | LD | 70297-L | EZZ14377TC-L | LOANER-TXCELL.LD.RETURN INTO INVENTORY | loaner return | test,clean and inspect. | BCTechSupport@iridex.com | 2025-04-22 06:08 | Yes | No |  | No |  | No
SRV-RQ002896 | SD | CUST-CYCLO-G6 | IQI93254C | CUST-G6, PSD ( $700.00 per ASM agreement), GUI upgrade | Upgrade software give to Robert Sandowal .JS | Upgraded to GUI and retested all the delivery devices as MFG spec.JS | BCTechSupport@iridex.com | 2025-04-22 07:12 | Yes | No |  | No |  | No
SRV-RQ002559 | PSD | CUST-SLX | 29157-LS | CUST-SLx, PSD, PM CHECUP | Low power out put. RY |  | BCTechSupport@iridex.com | 2025-04-22 10:13 | Yes | No |  | No |  | No
SRV-RQ002941 | PSD | CUST-IQ577 | IQY72300C | CUST-IQ577.PSD.E05111 | E05111.SD | No proplem found for the console. Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2025-04-22 10:40 | Yes | No |  | No |  | No
SRV-RQ002674 | PSI | 65600-08 | IQY72618C | CUST-IQ577 Topcon IQY72618C, E06002, PSI, provide estimate | Power output low, send laser to headroom for eval. SD | -Realigned the laser head, recalibrated the unit. | BCTechSupport@iridex.com | 2025-04-22 11:18 | Yes | No |  | No |  | No
SRV-RQ002655 | STG | 65777 | RC10320 | CUST-Remote.PSI. (Germany office) , RC10320, need to work with IQY72424C console | I HAVE NOT SEEN ANY DIFFERENCE BETWEEN CONSOLE AND REMOTE BOTH HAVE | RETESTED REMOTE IN HOUSE TOOL WORKS FINE.JS | BCTechSupport@iridex.com | 2025-04-22 12:46 | Yes | No |  | No |  | No
SRV-RQ002965 | LD | 66245-L | IQI91268C-L | LOANER-G6.LD.RETURN INTO INVENTORY | Loaner return.SD | Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2025-04-22 12:47 | Yes | No |  | No |  | No
SRV-RQ002800 | PSD | 32000-4 | TX19230C | CUST-TX, PSD, PM, need estimate ( $2800.00), if problems add labor and parts | ERROR 5. Low power at 2500mw .Laser head require Eval. CT |  | BCTechSupport@iridex.com | 2025-04-22 14:07 | Yes | No |  | No |  | No
SRV-RQ002885 | SD | CUST-IQ577 | IQY71626C | CUST-IQ577, PSD, LOW POWER | CUSTOMER COMPLAINT LOW POWER HAVE NOT DUPLICATED.JS | RETESTED ALL THE DELIVERY DEVICES AS MFG SPEC.JS | BCTechSupport@iridex.com | 2025-04-23 08:10 | Yes | No |  | No |  | No
SRV-RQ002977 | SD | CUST-SLX | 28388-LS | CUST-SLX.PSD. UNKNOWN FIBER ERROR | Low power output. Need to replace 2 laser diodes, aiming beam diode |  | BCTechSupport@iridex.com | 2025-04-23 11:15 | Yes | No |  | No |  | No
SRV-RQ002645 | PSD | 32000-1 | TX1386REVB | A | Low power at 2500mw and ENT high dac out of range. o.96mw | Realign laser head to manufacture TX specification. Yo | BCTechSupport@iridex.com | 2025-04-23 12:41 | Yes | No |  | No |  | No
SRV-RQ002955 | LD | 66245-L | IQI91107C-L | LOANER-G6.LD.RETURN INTO INVENTORY | Loaner return.SD | Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2025-04-23 14:15 | Yes | No |  | No |  | No
SRV-RQ002839 | SD | CUST-SL | 1675 | CUST-SL, PSD, OLD UNIT NEED PM DONE |  |  | BCTechSupport@iridex.com | 2025-04-24 06:01 | Yes | No |  | No |  | No
SRV-RQ002951 | LD | 66245-2-L | IQI92300C-L | LOANER-G6.LD.RETURN INTO INVENTORY | Loaner return .SD | Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2025-04-24 08:14 | Yes | No |  | No |  | No
SRV-RQ002950 | LD | 66245-L | IQI91248C-L | LOANER-G6.LD.RETURN INTO INVENTORY | Loaner return. SD | Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2025-04-24 10:14 | Yes | No |  | No |  | No
SRV-RQ002836 | WD | 32000-4 | TX19571C | CUST-TX.WD.JUMPY ENCODER BOARD, customer complain about intermitted shots | Power is not stable . intermittent jumping .Laser head require Eval.CT | Realign laser head to manufacture TX specification. Yo | BCTechSupport@iridex.com | 2025-04-24 12:29 | Yes | No |  | No |  | No
SRV-RQ002964 | LD | 66245-L | IQI91015C-L | LOANER-G6.LD.RETURN INTO INVENTORY | Loaner return.SD | Retested all the delivery return. SD | BCTechSupport@iridex.com | 2025-04-24 13:55 | Yes | No |  | No |  | No
SRV-RQ002673 | WI |  |  | CUST-IQ577 Topcon IQY72623C, bad port one, WI | Bad port one and power output low. Send laser to headroom for eval. SD-The LCM doesn't function properly, Low power | Replace resistance fiber detect board. SD-Replaced new LCM S/N 1JVAN2W from S/N 1JV8R1D, realigned the laser | BCTechSupport@iridex.com | 2025-04-24 16:07 | Yes | No |  | No |  | No
SRV-RQ002673 | WI | 65600-08 | IQY72623C | CUST-IQ577 Topcon IQY72623C, bad port one, WI | Bad port one and power output low. Send laser to headroom for eval. SD | Replace resistance fiber detect board. SD | BCTechSupport@iridex.com | 2025-04-24 16:07 | Yes | No |  | No |  | No
SRV-RQ002806 | PSD | 66245 | IQI91740C | CUST- CYCLO G6, PSD, PM | Unit came here for PM check Found UIM fail to load application because of encode |  | BCTechSupport@iridex.com | 2025-04-25 08:54 | Yes | No |  | No |  | No
SRV-RQ002497 | SCD | 65600-01 | IQY71159C | CUST-IQ577, SCD PP+, PM CHECKUP | UNIT NOT POWER UP FOUND BAD POWER SUPPLY. JS |  | BCTechSupport@iridex.com | 2025-04-25 09:58 | Yes | No |  | No |  | No
SRV-RQ002996 | PSD | CUST-LIO | 100684225 | CUST-LIO.PSD.FIBER REPLACEMENT | Cannot test due to Broken assebly fiber optic will need adjustment and alignment |  | BCTechSupport@iridex.com | 2025-04-28 11:32 | Yes | No |  | No |  | No
SRV-RQ003006 | SD | CUST-IQ577 | IQY72877C | CUST-IQ577, WD, BURNING SMELL, BEEPING A LOT | CUSTOMR COMPLAINT BURNING SMELL AND BEEPING, BECAUSE OF HAS NO TEMPRATURE | Replaced LCM PCBA FROM SN:1JVAN26 TO 1JVAN2U. Recalibration and test | BCTechSupport@iridex.com | 2025-04-28 12:22 | Yes | No |  | No |  | No
SRV-RQ002972 | SD | CUST-IQ532 | IQG81623C | CUST-IQ532,WD, NOT RECOGNIZING FIBERS, CHECK SHIP TO | CUSTOMER COMPLAINT NOT RECOGNISE FIBER. fOUND UIM NOT LOADING APPLICATION | REPLACED ENCODER LCD DISPLAY AND UIM AND RETESTED ALL THE DELIVERY DEVICES.JS | BCTechSupport@iridex.com | 2025-04-28 12:49 | Yes | No |  | No |  | No
SRV-RQ002923 | PSD | 11180-H | L2042 | CUST-LIO.PSD.PM | Cannot tes duet to Burnt assembly fiber optic, bent control stem, burnt halogen |  | BCTechSupport@iridex.com | 2025-04-28 13:52 | Yes | No |  | No |  | No
SRV-RQ003011 | SD | CUST-LIO | 1341754 | CUST-LIO, PSD, BROKEN CABLES , NEED ALIGNMENT, urgent repair | Cannot test due to Twisted and burnt assembly fiber optic, bent control stem, |  | BCTechSupport@iridex.com | 2025-04-29 08:04 | Yes | No |  | No |  | No
SRV-RQ002789 | PSD | 31215 | GL15084C30 | CUST-GLx, PSD, Incompatible Device" and "Use with 810mm Console | Laser is not recognize the fiber .Found fiber detect board pin1 came out. | Realign laser head to manufacture GL 3rd cooler specification. Yo | BCTechSupport@iridex.com | 2025-04-29 09:17 | Yes | No |  | No |  | No
SRV-RQ002940 | PSD | CUST-LIO | 1046538 | CUST-LIO, PSD, BROKEN KNOBS | Cannot test due to Twisted and burnt assembly fiber optic, broken control stem, |  | BCTechSupport@iridex.com | 2025-04-29 10:01 | Yes | No |  | No |  | No
SRV-RQ002914 | WD | 70295 | YH14522TC | Cust-TxCell,WD,E05111, please assess and repair/upgrade | CONTROL BOX NON-STOP BLINKING BAD HARD DRIVE. | REPLACED COMPUTER BOARD TESTED ALIGNED CALIBRATED CLEAN AND INSPECTED. | BCTechSupport@iridex.com | 2025-04-29 10:30 | Yes | No |  | No |  | No
SRV-RQ002966 | PSD | 30268-H | L2107 | CUST-LIO.PSD.PM | Cannot test due to Twisted and burnt assembly fiber optic, bent control stem, |  | BCTechSupport@iridex.com | 2025-04-29 11:35 | Yes | No |  | No |  | No
SRV-RQ002738 | PSI | CUST-IQ577 | IQY72792C | Cust-IQ577, WI, GUI problem , activtes by itself, replace the GUI | GUI problem , activtes by itself, replace the GUI.SD | Replace LCD display and retested all the delivery devices. SD | BCTechSupport@iridex.com | 2025-04-29 11:39 | Yes | No |  | No |  | No
SRV-RQ003005 | PSD | CUST-TXCELL | GH14685TC | CUST-TXCELL, psd, need service and estimate before completion | FILTER PLATE,GUIDE ADJUSTER,AND KNOB CLAMP ARE DEMAGE. |  | BCTechSupport@iridex.com | 2025-04-29 13:32 | Yes | No |  | No |  | No
SRV-RQ002905 | PSD | CUST-LIO | K2381 | Cust-LIO, PSD, mayneed fiberoptic replacment, estimte before completion | Cannot test due to Twiated and broken assembly fiber optic, missing eyepiece |  | BCTechSupport@iridex.com | 2025-04-29 14:00 | Yes | No |  | No |  | No
SRV-RQ002905 | PSD | CUST-LIO | K2381 | Cust-LIO, PSD, mayneed fiberoptic replacment, estimte before completion | Cannot test due to Twiated and broken assembly fiber optic, burnt halogen bulb |  | BCTechSupport@iridex.com | 2025-04-29 14:10 | Yes | No |  | No |  | No
SRV-RQ002273 | PSD | CUST-GL | GL13204B30-D | CUST-GL,PSD,NO ILLUMINATION, ADRESS CONFIRMED | No IIlumination . found LIO board #31208 burn .No part for repair . | No repair. Part 31208 is obsoleted. CT | BCTechSupport@iridex.com | 2025-04-29 14:22 | Yes | No |  | No |  | No
SRV-RQ002953 | DD | 66245D | IQI91662C-D | DEMO-G6.DD.RETURN INTO INVENTORY | Demo return. SD | Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2025-04-30 07:41 | Yes | No |  | No |  | No
SRV-RQ002990 | SD | CUST-IQ532 | IQG81812C | cust-iq532.WD.fiber optic board issue (PLATINUM) | FIBER NOT RECOGNISED AND FRONT BEZEL NOT CLOSED WITH CHASSIS. PAM NUT CAME OUT | Replaced fiber detect and Front bazel retested all the delivery devices.JS | BCTechSupport@iridex.com | 2025-04-30 08:30 | Yes | No |  | No |  | No
SRV-RQ003029 | LD | 66245-2-L | IQI92302C-L | Loaner. Cyclo G6 Laser. LD. Return to inventory | Loaner return. SD | Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2025-04-30 09:38 | Yes | No |  | No |  | No
SRV-RQ002939 | PSI | CUST-TXCELL | YEZ16062TC | PN:70292D , PSI , LED BLINKING GOES TO INVERNTORY | CONTROL BOX NON-STOP BLINKING, BAD COMPUTER BOARD. |  | BCTechSupport@iridex.com | 2025-04-30 10:18 | Yes | No |  | No |  | No
SRV-RQ002983 | WD | CUST-SLX | 30350-LS | CUST-SLx, WD , NO ILLUMINATION, VERIFY SHIP TO | No illumination light. RY | Replaced power supply PCBA from SN: 23030030 REV. M to SN: 2409-0041 | BCTechSupport@iridex.com | 2025-04-30 12:40 | Yes | No |  | No |  | No
SRV-RQ003046 | SI | 66245D-1 | IQI93771C-D | G6-DI.Return into inventory | Customer unit. Need to covert to demo and move it to FGI. | CONVERTED THE FOOT SWITCH TO DEMO AND RETESTED IT. JS | BCTechSupport@iridex.com | 2025-04-30 13:42 | Yes | No |  | No |  | No
SRV-RQ002987 | SD | CUST-IQ532 | IQG81333C | CUST-IQ532.PSD.E05111 | Customer complaint E05111 have not duplicated but found low power | -Replaced fine adjust screws, realigned the laser head, recalibrated the unit. | BCTechSupport@iridex.com | 2025-04-30 15:15 | Yes | No |  | No |  | No
SRV-RQ002929 | PSD | 30903-H500 | K2111 | CUST-LIO.PSD.PM | Cannot test due to Twisted and burnt assembly fiber optic, loose safety filter |  | BCTechSupport@iridex.com | 2025-04-30 15:54 | Yes | No |  | No |  | No
SRV-RQ002929 | PSD | 30903-H500 | K2111 | CUST-LIO.PSD.PM | Cannot test due to Twisted and burnt assembly fiber optic, loose safety filter |  | BCTechSupport@iridex.com | 2025-04-30 16:06 | Yes | No |  | No |  | No
SRV-RQ002937 | PSD | CUST-CYCLO-G6 | IQI91129C | CUST- G6, PSD, PM CHECKUP | Customer return for PM check .SD | Retested all the delivery device.SD | BCTechSupport@iridex.com | 2025-05-01 07:40 | Yes | No |  | No |  | No
SRV-RQ002999 | PSD | CUST-LIO | K1007 | CUST-LIO.SCD. FIBER DAMAGED | Cannot test duet to Twiated and broken assembly fiber optic, burnt halogen |  | BCTechSupport@iridex.com | 2025-05-01 10:34 | Yes | No |  | No |  | No
SRV-RQ002928 | PSD | CUST-LIO | L1947 | CUST-LIO, PSD, BROKEN, NO PWER | Cannot test duet to Twisted and broken assembly fiber optic, bent control stem, |  | BCTechSupport@iridex.com | 2025-05-01 12:13 | Yes | No |  | No |  | No
SRV-RQ002713 | PSI | 33004-1 | GL41204C | Cust-GL. PSI. (2015), need power adjustment since is lower end or out of spec output. | Display bad #77022 ,the letters are jumbled and unclear .Error 5 keep reseting. | Realign laser head to manufacture Neo Specification. Yo | BCTechSupport@iridex.com | 2025-05-01 13:01 | Yes | No |  | No |  | No
SRV-RQ002971 | WI | 65600-08 | IQY72739C | CUST-IQ577 WI, OBF "Wavelengths Incompatible device.", the panel beeps on its own | OBF "Wavelengths Incompatible device.", the panel beeps on its own.sd | No problem found. restested all the delivery devices. sd | BCTechSupport@iridex.com | 2025-05-01 13:11 | Yes | No |  | No |  | No
SRV-RQ003000 | PSD | CUST-CYCLO-G6 | IQI92347C | CUST-CYCLO-G6, PSD , PM CHECKUP | Customer return for PM check. SD | Restested all the delivery devices. SD | BCTechSupport@iridex.com | 2025-05-01 14:20 | Yes | No |  | No |  | No
SRV-RQ002879 | PSD | CUST-LIO | 1361977 | CUST-LIO, PSD, BRKEN NEED ALIGNMENT | Cannot test due to Twisted and burnt assembly fiber optic, loose housing |  | BCTechSupport@iridex.com | 2025-05-01 14:51 | Yes | No |  | No |  | No
SRV-RQ003009 | SD | CUST-CYCLO-G6 | IQI93152C | CUST-G6.PSD.PM | PM CHECK | RETESTED ALL THE DELIVERY DEVICES.JS | BCTechSupport@iridex.com | 2025-05-02 08:53 | Yes | No |  | No |  | No
SRV-RQ003018 | PSD | CUST-LIO | 1039166 | CUST-LIO.PSD.PM | Cannot test due to Twisted and burnt assembly fiber optic, burnt halogen bulb, |  | BCTechSupport@iridex.com | 2025-05-02 10:35 | Yes | No |  | No |  | No
SRV-RQ002877 | PSD | CUST-TXCELL | YEZ14173TC | CUST-TxCell, PSD, LED KEPT BLINKING | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING BAD COMPUTER BOARD. |  | BCTechSupport@iridex.com | 2025-05-02 10:51 | Yes | No |  | No |  | No
SRV-RQ002992 | PSD | CUST-CYCLO-G6 | IQI91550C | CUST-CYCLO-G6, PSD, JUST REPAIRED ,NO AUDIBLE, MAX LEVEL WHEN FOOTSWITCH DEPRESS, VERIFY SHIP TO | No speaker sound. SD | Replace assembly cable speaker, and retested all the delivery devices. SD | BCTechSupport@iridex.com | 2025-05-02 12:01 | Yes | No |  | No |  | No
SRV-RQ003016 | SD | CUST-IQ532 | IQG82171C | CUST-IQ532, WD , NO OUTPUT MAY BE POWER SUPPLY BAD, CALL BEFORE SHIPPING | CUSTOMER COMPLAINT NO POWER OUTPUT,PROBLEM NOT DUPLICATED.JS | RETESTED ALL THE DELIVERY DEVICES AS MFG SPEC.JS | BCTechSupport@iridex.com | 2025-05-02 12:28 | Yes | No |  | No |  | No
SRV-RQ002489 | PSD | CUST-LIO | K0793 | CUST-LIO.PSD.KNOBS NOT MOVING |  | Cannot test due to Twisted and burnt assebly fiber optic, burnt halogen bulb and | BCTechSupport@iridex.com | 2025-05-02 12:57 | Yes | No |  | No |  | No
SRV-RQ002976 | PSD | CUST-LIO | K1543 | cust-lio.psd,broken mrors | Cannot tes due to Twisted and burnt assembly fiber optic, burnt halogen bulb, |  | BCTechSupport@iridex.com | 2025-05-02 14:21 | Yes | No |  | No |  | No
SRV-RQ002982 | PSD | CUST-TXCELL | YEZ14515TC | CUST-TxCell, PSD, LED KEEP BLINKING | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING |  | BCTechSupport@iridex.com | 2025-05-05 06:40 | Yes | No |  | No |  | No
SRV-RQ002933 | PSD | CUST-FOOTSWITCH | 110907F | CUST-FOOTSWITCH, PSD, DEAD BATTRIES | Footswitch is not working . No communication. | No communication . Need to send back to supplier for repair.CT | BCTechSupport@iridex.com | 2025-05-05 06:50 | Yes | No |  | No |  | No
SRV-RQ002862 | WD | 31808 | 112556F | PN:31808 , RETURND TO INVENTORY,WD,REPLACMENT SENT 950+S&h | Return to inventory. Need Convert footswitch to Demo. CT | Retested and Converted footswitch from 31808 to 31808D. CT | BCTechSupport@iridex.com | 2025-05-05 07:17 | Yes | No |  | No |  | No
SRV-RQ002995 | PSD | CUST-LIO | 1362284 | CUST-LIO. PSD. BROKEN MIRROR | Cannot test due to loose mirror turn from cartridge, Twisted and broken |  | BCTechSupport@iridex.com | 2025-05-05 08:49 | Yes | No |  | No |  | No
SRV-RQ003034 | LD | 13030-L | 29553-LS-L | Loaners. SLx. LD. Loaner return to inventory | Loaner returned. RY | Retested to MFG test traveler spec. RY | BCTechSupport@iridex.com | 2025-05-05 09:47 | Yes | No |  | No |  | No
SRV-RQ002739 | PSD | 70292 | YEZ14144TC | CUST-TXCELL, PSD, STATES THAT TXCELL CTRL BOX WILL NOT LOAD UP/ GREEN LIGHT DOESN'T STAY SOLID | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING |  | BCTechSupport@iridex.com | 2025-05-05 10:07 | Yes | No |  | No |  | No
SRV-RQ003014 | PSD | CUST-LIO | 1265998 | CUST-LIO.PSD.PM | Cannote test due to Twisted and burnt assembly fiber optic, broken control stem, |  | BCTechSupport@iridex.com | 2025-05-05 10:27 | Yes | No |  | No |  | No
SRV-RQ003028 | SD | CUST-SLX | 29735-LS | cust-SLX.PSD.Not decting fiber optic cables |  |  | BCTechSupport@iridex.com | 2025-05-05 10:53 | Yes | No |  | No |  | No
SRV-RQ002832 | WD | CUST-LIO | 1358671 | CUST, LIO, WD, HAS BEEN DROPPED AND BROKEN | Cannot test due to Broken control stem, Twisted and burnt assembly fiber optic, |  | BCTechSupport@iridex.com | 2025-05-05 12:35 | Yes | No |  | No |  | No
SRV-RQ002948 | LD | 65500-L | IQG81290C-L | LOANER-IQ532.LD.RETURN INTO INVENTORY (CAME FROM 24910) | Loaner return. SD | Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2025-05-05 13:22 | Yes | No |  | No |  | No
SRV-RQ003001 | WD | CUST-TXCELL | YH14111TC | cust-txcell.WD. damaged when it was shipped | CONTROL BOX NON-STOP BLINKING. BEZEL AND ASSY, CABLE AC ARE DAMAGE | REPLACED COMPUTER BOARD TESTED ALIGNED CALIBRATED CLEAN AND INSPECTED. | BCTechSupport@iridex.com | 2025-05-06 07:22 | Yes | No |  | No |  | No
SRV-RQ002924 | SD | CUST-TXCELL | EZZ13301TC | CUST-TxCell, PSD, LED BLINKING , AUX REQD MESSAGE, VERIFY SHIP TO | UNABLE TO TESTED CONTROL BOX NON-STOP BLINKING COMPUTER BOARD NOT WORKING |  | BCTechSupport@iridex.com | 2025-05-06 08:44 | Yes | No |  | No |  | No
SRV-RQ002903 | SI | CUST-IQ577 | IQY71444C | Cust-IQ577, E06010, PSI, provide estimate before completion | EO6010 CUSTOMETR COMPLAT AND DUPLICATED AS WELL,HEAD SENT TO HEAD ROOM.JS |  | BCTechSupport@iridex.com | 2025-05-06 09:04 | Yes | No |  | No |  | No
SRV-RQ003050 | PSD | CUST-LIO | K1478 | CUST-LIO, PSD, BROKEN, LOW POWER | Cannot test due to Twisted and burnt assembly fiber optic, missing eyepiece |  | BCTechSupport@iridex.com | 2025-05-06 10:39 | Yes | No |  | No |  | No
SRV-RQ003073 | PSD | CUST-LIO | 1143461 | CUST-LIO. PSD. PM | Cannot test due to Twisted and broken assembly fiber optic, burnt halogen, |  | BCTechSupport@iridex.com | 2025-05-06 10:59 | Yes | No |  | No |  | No
SRV-RQ002946 | PSD | CUST-SLA | GH6398 | CUST-SLA, PSD, NO AIM BEAM, NEED ALIGNMENT, | Broken fiber, | Repalced fiber, Test, clean and inspect. | BCTechSupport@iridex.com | 2025-05-06 14:05 | Yes | No |  | No |  | No
SRV-RQ003025 | CEM | CUST-GLX | GLX18828C | CUST-GLX.CEM.PM | In house tool for PM check up .No problem found.CT | Recalibrated all delivery devices ,retest and burn in ok. CT | BCTechSupport@iridex.com | 2025-05-06 14:26 | Yes | No |  | No |  | No
SRV-RQ002978 | WD | CUST-CYCLO-G6 | IQI92680C | CUST-G6.WD.PM | Customer return for PM check. SD | Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2025-05-07 10:46 | Yes | No |  | No |  | No
SRV-RQ003003 | PSD | CUST-SLA | GH9429LM | CUST-SLA.PSD.PM |  |  | BCTechSupport@iridex.com | 2025-05-07 10:51 | Yes | No |  | No |  | No
SRV-RQ002803 | PSD | 32000-1 | TX18662C | CUST-TX.PSD.POWER SUPPLY ISSUE | Errror 2 and Error 20 keep reset. Clicking noise found CTRL board #31586-1 is | Realign laser head to manufacture TX specification. Yo | BCTechSupport@iridex.com | 2025-05-07 12:27 | Yes | No |  | No |  | No
SRV-RQ002728 | PSI | 32000-4 | TX19414C | LASER REPORTED AN ERROR OF LOWER POWER/ LASER CAVITY/ NO ENERGY | Error 5 .Power shut off when firing . Laser head require Eval. CT |  | BCTechSupport@iridex.com | 2025-05-07 13:31 | Yes | No |  | No |  | No
SRV-RQ003043 | WD | CUST-LIO | K2177 | CUST- WD, LIO, oculars seem mildly misaligned, headband loosens very easily, without turning the kno | Cannot test due to loose housing illumination, stripped aiming beam limit screw, | Replaced eyepiece rubber, aiming beam limit screw, adjusted housing illumination | BCTechSupport@iridex.com | 2025-05-07 14:31 | Yes | No |  | No |  | No
SRV-RQ003022 | SD | CUST-CYCLO-G6 | IQI91050C | CUST-CYCLO-G6.CEM.PM | LOW POWER OUTPUT JS | Replaced fiber adapter. Realign and calibration laser head. RY | BCTechSupport@iridex.com | 2025-05-08 08:54 | Yes | No |  | No |  | No
SRV-RQ002932 | SD | CUST-IQ577 | IQY72564C | CUST-IQ577, PSD, UNABLE TO SCREW IN TIGHT SLA & LIO, VERIFY SHIPTO | FIBER PORT WAS MIALIGNED,AFTER FIXED FOUND LOW POWER,HR.JS | -Realigned the laser head, recalibrated the unit. | BCTechSupport@iridex.com | 2025-05-08 10:03 | Yes | No |  | No |  | No
SRV-RQ003040 | PSD | CUST-LIO | 1344105 | CUST-LIO, PSD, LOW POWER, BROKEN FIBER, VERIFY SHIP TO | Cannot test due to Twisted and broken assembly fiber optic, bent control stem, |  | BCTechSupport@iridex.com | 2025-05-08 12:00 | Yes | No |  | No |  | No
SRV-RQ003054 | PSD | CUST-LIO | L3155 | CUST-LIO, PSD, UNABLE TO CENTER THE IMAGE, VERY SHIP TO ADRESS | Cannot test due to Twisted and burt assembly fiber optic, wrong and burnt |  | BCTechSupport@iridex.com | 2025-05-08 12:58 | Yes | No |  | No |  | No
SRV-RQ002944 | WD | CUST-TXCELL | EZZ14071TC | CUST-TXCELL. WD .LED BLINKNING (HOT) | CONTROL BOX NON-STOP BLINKING COMPUTER BOARD NOT WORKING | REPLACED COMPUTER BOARD BEZEL.ELEVETOR CUP,BALL PIN. | BCTechSupport@iridex.com | 2025-05-08 13:10 | Yes | No |  | No |  | No
SRV-RQ002617 | PSI | CUST-SLA | GH16675LM | Cust-SLA PSI, is not detected by a GLx ( check and verify its PCB operations) | Broken fiber. | Repalced fiber,Test,clean and inspec. | BCTechSupport@iridex.com | 2025-05-09 07:32 | Yes | No |  | No |  | No
SRV-RQ003071 | DCI | CUST-IQ577 | IQY71570C | DEMO, DCI, Clinical study use, contact Joan Stuffer for shipping out to Carlton | Customer return for calibration due. SD | Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2025-05-12 06:54 | Yes | No |  | No |  | No
SRV-RQ003078 | SD | CUST-IQ577 | IQY71942C | CUST-IQ577. PSD. the display screen is not lighting up | DAMAGED FRONT BEZEL AND NOT POWER UP | REPLACED FRONT BEZEL,SCM AND BRAND IQ LABEL AND RETESTED ALL | BCTechSupport@iridex.com | 2025-05-12 06:59 | Yes | No |  | No |  | No
SRV-RQ002804 | PSD | CUST-TX | TX18100C | CUST-TX.PSD.POWER SUPPLY ISSUE | Reuse all original optics. Yo |  | BCTechSupport@iridex.com | 2025-05-13 08:28 | Yes | No |  | No |  | No
SRV-RQ003076 | PSD | CUST-TXCELL | YEZ13662TC | CUST-TXCELL, PSD, LED KEEP BLINKING | CONTROL BOX NON-STOP BLINKING, COMPUTER BOARD NOT WORKING. |  | BCTechSupport@iridex.com | 2025-05-13 15:19 | Yes | No |  | No |  | No
SRV-RQ003074 | SD | CUST-SLX | 29215-LS | CUST-SLx. PSD. PM (ERROR 15) | LOW POWER AND PHOTOCELL OUT OF RANGE HEAD SENT TO HEAD ROOM.JS | REPLACED AIMING BEAM DIODE, ADJUST SCREWS AND REALIGN LASER | BCTechSupport@iridex.com | 2025-05-14 12:38 | Yes | No |  | No |  | No
SRV-RQ002985 | PSD | CUST-TX | TX1439REVB | CUST-TX, PSD, ILLUMINATION RCA CONNECTION IS LOOSE, NO ILLUMINATION | LIO port broken # 31900. Error 5 .power shuts off when firing. | Reused all optics and realigned laser head to mfg TX spec. YO | BCTechSupport@iridex.com | 2025-05-15 08:16 | Yes | No |  | No |  | No
SRV-RQ003041 | PSD | CUST-CYCLO-G6 | IQI91120C | CUST-G6,PSD, PM | Customer return for PM check. SD | Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2025-05-15 12:57 | Yes | No |  | No |  | No
SRV-RQ002863 | SD | CUST-IQ577 | IQY71747C | Cust-IQ577, PSD, E05111: Delivery device or SmarkKey/AuxDevice not connected. | E05111 DELIVERY DEVICES NOT RECOGNIZED. JS | NO PROBLEM FOUND, UNIT WORKS FINE. JS | BCTechSupport@iridex.com | 2025-05-16 07:45 | Yes | No |  | No |  | No
SRV-RQ002997 | SD | CUST-IQ532 | IQG82044C | CUST-IQ532.PSD.UNIT IS BEEPING CONSTANT | NO BOOTING UP FAN START FUUL BLOWN HEAD SENT TO HEAD ROOM.JS | -Replaced new LCM S/N 1JVAP0A from S/N 1JV8R0 realigned the laser head, | BCTechSupport@iridex.com | 2025-05-16 11:06 | Yes | No |  | No |  | No
SRV-RQ002832 | WD | CUST-LIO | 1358671 | CUST, LIO, WD, HAS BEEN DROPPED AND BROKEN | Cannot test due to Broken control stem, viewing assembly dislocated from T-slot | Replaced control stem, eccentric shaft, adjusted T-slot adapter, adjusted | BCTechSupport@iridex.com | 2025-05-16 14:33 | Yes | No |  | No |  | No
SRV-RQ003085 | PSD | CUST-LIO | K1924 | CUST-LIO, PSD, NO ILLUMINATION NEED ALIGNMENT | Cannote test due to Burnt assembly fiber optic, missing eyepiece rubber, loose | Replaced eyepiece rubber, dust cap with lanyard, halogen bulb, adjusted T-slot | BCTechSupport@iridex.com | 2025-05-19 12:48 | Yes | No |  | No |  | No
SRV-RQ003051 | SD | CUST-SLX | 30146-LS | CUST-SLx, PSD, NO ILLUMINATION |  |  | BCTechSupport@iridex.com | 2025-05-20 06:45 | Yes | No |  | No |  | No
SRV-RQ003012 | PSD |  |  | CUST-IQ577, PSD, PORT ONE NOT RECOGNIZING FIBER, VERIFY SHIP TO | PORT ONE NOT RECOGBIZING FIBER. SDPORT ONE NOT RECOGNIZING FIBER.SD | No problem found. Retested all the delivery divices. SDNo prolem found. Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2025-05-20 10:52 | Yes | No |  | No |  | No
SRV-RQ003012 | PSD | CUST-IQ577 | IQY71253C-D | CUST-IQ577, PSD, PORT ONE NOT RECOGNIZING FIBER, VERIFY SHIP TO | PORT ONE NOT RECOGBIZING FIBER. SDPORT ONE NOT RECOGNIZING FIBER.SD | No problem found. Retested all the delivery divices. SDNo prolem found. Retested all the delivery devices. SD | BCTechSupport@iridex.com | 2025-05-20 10:52 | Yes | No |  | No |  | No
SRV-RQ002629 | SI | 13030 | 30487-LS | SLx console, WI, reported as OBF, error 1. WI. Deliver to Micheal, Mfg. Eng. only |  |  | BCTechSupport@iridex.com | 2025-05-20 12:44 | Yes | No |  | No |  | No